## BC SNR
In-built Body Coil SNR
| Transverse | Sagittal | Coronal | Mean
Signal to noise ratio | 47.7776 | 48.5835 | 50.2148 | 48.8587
Normalised SNR | 448.8578 | 456.4287 | 471.7549 | 459.0138
Plane: Transverse |  |  |  | Plane: Sagittal |  |  | Plane: Coronal
ROI | Signal Mean | Difference Standard Deviation |  | ROI | Signal Mean | Difference Standard Deviation | ROI | Signal Mean | Difference Standard Deviation
1 | 1820.02 | 55.519 |  | 1 | 1874.533 | 58.187 | 1 | 1886.595 | 56.437
2 | 1870.755 | 63.032 |  | 2 | 1924.55 | 48.661 | 2 | 1882.862 | 54.343
3 | 1935.23 | 57.445 |  | 3 | 1937.67 | 63.642 | 3 | 1955.015 | 59.613
4 | 1946.348 | 52.946 |  | 4 | 1837.027 | 54.107 | 4 | 1867.86 | 46.982
5 | 1846.085 | 49.843 |  | 5 | 1902.52 | 51.248 | 5 | 1876.957 | 49.311
Signal to noise ratio | 47.7776 |  |  | Signal to noise ratio | 48.5835 |  | Signal to noise ratio | 50.2148
Input Parameters | Value
Phase encode matrix size | 256 | [5 empty] | BC_TRA_SNR_Images_with_ROIs
Frequency encode matrix size | 256
Repetition time (ms) | 1000
Number of signal averages | 1
Field-of-view (mm) | 250
Bandwidth from scanner interface | 130
Scanner manufacturer | Siemens
Static magnetic field strength (T) | 1.5
Q-factor for head or body | 1
Q-factor for test object | 1
Calculated Parameters | Value
±Bandwidth (kHz) | 16.64
Voxel size in the frequency encode direction (mm) | 0.97656
Voxel size in the phase encode direction (mm) | 0.97656
Slice width for transverse acquisition (mm) | 5.1953 | [5 empty] | BC_SAG_SNR_Images_with_ROIs
Slice width for sagittal acquisition (mm) | 5.1953
Slice width for coronal acquisition (mm) | 5.1953
Fixed Parameters | Value
±Nominal bandwidth (kHz) | 30
SNR Normalisation Factors | Value
Bandwidth normalisation | 0.74476
Voxel size correction (transverse) | 201.8312
Voxel size correction (sagittal) | 201.8312
Voxel size correction (coronal) | 201.8312
Scan time correction | 0.0625
Coil loading normalisation | 1
BC_COR_SNR_Images_with_ROIs
BW_siemens_ |  | #NAME?
BW_ge_ |  | #NAME?
BW_philips_
Note on A22: Wastling,Stephen: Readout bandwidth from scanner: - GE in +/-kHz (typically around ±16 kHz) - Philips fat-water shift in Px (typically 1.69 at 1.5T and 3.38 at 3T) - Siemens in Hz/Px (typically 130 Hz/Px)
Note on A23: Wastling,Stephen: Scanner manufacturer, used when calculating the bandwidth in ±kHz
Note on A24: Wastling,Stephen: Used when calculating the bandwidth in ±kHz from the fat-water shift for Philips scanners
Note on A25: Wastling,Stephen: Q-factor for the head (typically set to 1)
Note on A26: Wastling,Stephen: Q-factor for the test object (typically set to 1)
Note on A32: Wastling,Stephen: Copied from slice width sheet
Note on A33: Wastling,Stephen: Copied from slice width sheet
Note on A34: Wastling,Stephen: Copied from slice width sheet
Note on A37: Wastling,Stephen: ± Nominal bandwidth (kHz). Set to ± 30 kHz
Note on A40: Wastling,Stephen: Bandwidth normalisation
Note on A41: Wastling,Stephen: Voxel size correction. As per MagNET guidance, voxel sizes are converted to cm.
Note on A44: Wastling,Stephen: Scan-time correction. As per MagNET guidance, TR is converted to seconds
Note on A45: Wastling,Stephen: Coil loading normalisation
Instructions:
1. Enter the acquisition parameters into cells B17:B26
2. Copy the values from BC_TRA_SNR.csv into cells B9:C13
3. Copy the values from BC_SAG_SNR.csv into cells F9:G13
4. Copy the values from BC_COR_SNR.csv into cells J9:K13
5. Copy BC_TRA_SNR_Image_with_ROIs.png over cell I19
6. Copy BC_SAG_SNR_Image_with_ROIs.png over cell I33
7. Copy BC_COR_SNR_Image_with_ROIs.png over cell I48
The normailsed SNR (NSNR) is given by:
with measured in seconds, and the voxel sizes ( and ) in cm.
The nominal bandwith is and and are set to 1.
The normailsed SNR (NSNR) is given by:
𝑁𝑆𝑁𝑅=√(𝐵𝑊/(𝐵𝑊_𝑛𝑜𝑚 ).) 1/Δ𝑥Δ𝑦Δ𝑧.1/√(𝑁_𝑝𝑒.𝑇𝑅.𝑁𝑆𝐴).√(𝑄_ℎ𝑒𝑎𝑑/𝑄_(𝑡𝑒𝑠𝑡−𝑜𝑏𝑗) ).𝑆𝑁𝑅
with 𝑇𝑅 measured in seconds, and the voxel sizes (Δ𝑥,Δ𝑦 and Δ𝑧) in cm.
The nominal bandwith is ±30 𝑘𝐻𝑧 and 𝑄_ℎ𝑒𝑎𝑑 and 𝑄_(𝑡𝑒𝑠𝑡−𝑜𝑏𝑗) are set to 1.

## HC SNR
Head Coil SNR
| Transverse | Sagittal | Coronal | Mean
Signal to noise ratio | 377.6806 | 278.6668 | 310.2707 | 322.206
Normalised SNR | 3548.2066 | 2617.9988 | 2914.9089 | 3027.0381
Plane: Transverse |  |  |  | Plane: Sagittal |  |  | Plane: Coronal
ROI | Signal Mean | Difference Standard Deviation |  | ROI | Signal Mean | Difference Standard Deviation | ROI | Signal Mean | Difference Standard Deviation
1 | 1402.608 | 4.755 |  | 1 | 1167.65 | 6.321 | 1 | 1354.3 | 5.066
2 | 1434.078 | 5.256 |  | 2 | 1340.243 | 5.387 | 2 | 1367.105 | 5.603
3 | 895.045 | 6.983 |  | 3 | 886.612 | 6.609 | 3 | 894.795 | 7.516
4 | 1561.025 | 4.532 |  | 4 | 911.747 | 5.083 | 4 | 1171.572 | 4.448
5 | 1626.247 | 4.382 |  | 5 | 1267.818 | 4.888 | 5 | 1166.375 | 4.506
Signal to noise ratio | 377.6806 |  |  | Signal to noise ratio | 278.6668 |  | Signal to noise ratio | 310.2707
Input Parameters | Value
Phase encode matrix size | 256 | [5 empty] | HC_TRA_SNR_Images_with_ROIs
Frequency encode matrix size | 256
Repetition time (ms) | 1000
Number of signal averages | 1
Field-of-view (mm) | 250
Bandwidth from scanner interface | 130
Scanner manufacturer | Siemens
Static magnetic field strength (T) | 1.5
Q-factor for head or body | 1
Q-factor for test object | 1
Calculated Parameters | Value
±Bandwidth (kHz) | 16.64
Voxel size in the frequency encode direction (mm) | 0.97656
Voxel size in the phase encode direction (mm) | 0.97656
Slice width for transverse acquisition (mm) | 5.1953 | [5 empty] | HC_SAG_SNR_Images_with_ROIs
Slice width for sagittal acquisition (mm) | 5.1953
Slice width for coronal acquisition (mm) | 5.1953
Fixed Parameters | Value
±Nominal bandwidth (kHz) | 30
SNR Normalisation Factors | Value
Bandwidth normalisation | 0.74476
Voxel size correction (transverse) | 201.8312
Voxel size correction (sagittal) | 201.8312
Voxel size correction (coronal) | 201.8312
Scan time correction | 0.0625
Coil loading normalisation | 1
HC_COR_SNR_Images_with_ROIs
BW_siemens_ |  | #NAME?
BW_ge_ |  | #NAME?
BW_philips_
Note on A22: Wastling,Stephen: Readout bandwidth from scanner: - GE in +/-kHz (typically around ±16 kHz) - Philips fat-water shift in Px (typically 1.69 at 1.5T and 3.38 at 3T) - Siemens in Hz/Px (typically 130 Hz/Px)
Note on A23: Wastling,Stephen: Scanner manufacturer, used when calculating the bandwidth in ±kHz
Note on A24: Wastling,Stephen: Used when calculating the bandwidth in ±kHz from the fat-water shift for Philips scanners
Note on A25: Wastling,Stephen: Q-factor for the head (typically set to 1)
Note on A26: Wastling,Stephen: Q-factor for the test object (typically set to 1)
Note on A32: Wastling,Stephen: Copied from slice width sheet
Note on A33: Wastling,Stephen: Copied from slice width sheet
Note on A34: Wastling,Stephen: Copied from slice width sheet
Note on A37: Wastling,Stephen: ± Nominal bandwidth (kHz). Set to ± 30 kHz
Note on A40: Wastling,Stephen: Bandwidth normalisation
Note on A41: Wastling,Stephen: Voxel size correction. As per MagNET guidance, voxel sizes are converted to cm.
Note on A44: Wastling,Stephen: Scan-time correction. As per MagNET guidance, TR is converted to seconds
Note on A45: Wastling,Stephen: Coil loading normalisation
The normailsed SNR (NSNR) is given by:
with measured in seconds, and the voxel sizes ( and ) in cm.
The nominal bandwith is and and are set to 1.
The normailsed SNR (NSNR) is given by:
𝑁𝑆𝑁𝑅=√(𝐵𝑊/(𝐵𝑊_𝑛𝑜𝑚 ).) 1/Δ𝑥Δ𝑦Δ𝑧.1/√(𝑁_𝑝𝑒.𝑇𝑅.𝑁𝑆𝐴).√(𝑄_ℎ𝑒𝑎𝑑/𝑄_(𝑡𝑒𝑠𝑡−𝑜𝑏𝑗) ).𝑆𝑁𝑅
with 𝑇𝑅 measured in seconds, and the voxel sizes (Δ𝑥,Δ𝑦 and Δ𝑧) in cm.
The nominal bandwith is ±30 𝑘𝐻𝑧 and 𝑄_ℎ𝑒𝑎𝑑 and 𝑄_(𝑡𝑒𝑠𝑡−𝑜𝑏𝑗) are set to 1.
Instructions:
1. Enter the acquisition parameters into cells B17:B26
2. Copy the values from HC_TRA_SNR.csv into cells B9:C13
3. Copy the values from HC_SAG_SNR.csv into cells F9:G13
4. Copy the values from HC_COR_SNR.csv into cells J9:K13
5. Copy HC_TRA_SNR_Image_with_ROIs.png over cell I19
6. Copy HC_SAG_SNR_Image_with_ROIs.png over cell I33
7. Copy HC_COR_SNR_Image_with_ROIs.png over cell I48

## BC Uniformity
In-built Body Coil Uniformity
x-direction | y-direction | z-direction | Average | Standard Deviation
1 | 0.97813 | 0.90938 | 0.9625 | 0.043839
| Transverse | Sagittal | Coronal
Horizontal | 1 | 0.9875 | 1
Vertical | 0.96875 | 0.925 | 0.89375
A
P
R
L
H
F
A
P
H
F
R
L
TRA
SAG
COR
X
Y
Y
Z
X
Z
Instructions:
1. Follow the instructions included in sheets "BC Tra Uniformity", "BC Sag Uniformity" and "BC Cor Uniformity"
2. Copy the Uniformity Profile chart into the report
The x-direction is left-right, the y-direction is up-down and the z-direction is along the bore:
y
z
x

## BC Tra Uniformity
| Horizontal | Vertical
Fractional Uniformity | 1 | 0.96875
Horizontal |  |  |  | Vertical
Pixel | Profile Data | Central ROI Mean | Abs Difference | Pixel | Profile Data | Central ROI Mean | Abs Difference
1 | 1775.1 | 1935.23 | 160.13 | 1 | 1690 | 1935.23 | 245.23
2 | 1782.1 | 1935.23 | 153.13 | 2 | 1720.3 | 1935.23 | 214.93
3 | 1798.5 | 1935.23 | 136.73 | 3 | 1737.8 | 1935.23 | 197.43
4 | 1817.5 | 1935.23 | 117.73 | 4 | 1744.4 | 1935.23 | 190.83
5 | 1830.7 | 1935.23 | 104.53 | 5 | 1697.4 | 1935.23 | 237.83
6 | 1838.2 | 1935.23 | 97.03 | 6 | 1747.2 | 1935.23 | 188.03 | BC_TRA_UNIFORMITY_ROIs
7 | 1836.8 | 1935.23 | 98.43 | 7 | 1760.3 | 1935.23 | 174.93
8 | 1825.7 | 1935.23 | 109.53 | 8 | 1741.3 | 1935.23 | 193.93
9 | 1835 | 1935.23 | 100.23 | 9 | 1768.2 | 1935.23 | 167.03
10 | 1857.3 | 1935.23 | 77.93 | 10 | 1780.8 | 1935.23 | 154.43
11 | 1854.3 | 1935.23 | 80.93 | 11 | 1762.1 | 1935.23 | 173.13
12 | 1841.2 | 1935.23 | 94.03 | 12 | 1778 | 1935.23 | 157.23
13 | 1848.3 | 1935.23 | 86.93 | 13 | 1816.1 | 1935.23 | 119.13
14 | 1850.1 | 1935.23 | 85.13 | 14 | 1838.4 | 1935.23 | 96.83
15 | 1846.7 | 1935.23 | 88.53 | 15 | 1819.2 | 1935.23 | 116.03
16 | 1846.9 | 1935.23 | 88.33 | 16 | 1837.5 | 1935.23 | 97.73
17 | 1866.9 | 1935.23 | 68.33 | 17 | 1828.4 | 1935.23 | 106.83
18 | 1873.8 | 1935.23 | 61.43 | 18 | 1844.9 | 1935.23 | 90.33
19 | 1864.2 | 1935.23 | 71.03 | 19 | 1853.4 | 1935.23 | 81.83
20 | 1886.7 | 1935.23 | 48.53 | 20 | 1848.2 | 1935.23 | 87.03
21 | 1877.4 | 1935.23 | 57.83 | 21 | 1861.6 | 1935.23 | 73.63
22 | 1883.8 | 1935.23 | 51.43 | 22 | 1865.9 | 1935.23 | 69.33
23 | 1901 | 1935.23 | 34.23 | 23 | 1872.1 | 1935.23 | 63.13
24 | 1872.4 | 1935.23 | 62.83 | 24 | 1881.4 | 1935.23 | 53.83
25 | 1885.4 | 1935.23 | 49.83 | 25 | 1871.6 | 1935.23 | 63.63
26 | 1901.8 | 1935.23 | 33.43 | 26 | 1892.4 | 1935.23 | 42.83
27 | 1871.9 | 1935.23 | 63.33 | 27 | 1889 | 1935.23 | 46.23
28 | 1882.8 | 1935.23 | 52.43 | 28 | 1891.4 | 1935.23 | 43.83
29 | 1914.5 | 1935.23 | 20.73 | 29 | 1912.8 | 1935.23 | 22.43
30 | 1903.4 | 1935.23 | 31.83 | 30 | 1904.8 | 1935.23 | 30.43
31 | 1914.8 | 1935.23 | 20.43 | 31 | 1900.5 | 1935.23 | 34.73
32 | 1932.5 | 1935.23 | 2.73 | 32 | 1901.8 | 1935.23 | 33.43
33 | 1905.7 | 1935.23 | 29.53 | 33 | 1911.9 | 1935.23 | 23.33
34 | 1924.5 | 1935.23 | 10.73 | 34 | 1910.5 | 1935.23 | 24.73
35 | 1903.6 | 1935.23 | 31.63 | 35 | 1885.4 | 1935.23 | 49.83
36 | 1909.7 | 1935.23 | 25.53 | 36 | 1917.2 | 1935.23 | 18.03
37 | 1909.6 | 1935.23 | 25.63 | 37 | 1925.8 | 1935.23 | 9.43
38 | 1897.7 | 1935.23 | 37.53 | 38 | 1912.6 | 1935.23 | 22.63
39 | 1939.9 | 1935.23 | 4.67 | 39 | 1939.6 | 1935.23 | 4.37
40 | 1929.3 | 1935.23 | 5.93 | 40 | 1920.6 | 1935.23 | 14.63
41 | 1941.6 | 1935.23 | 6.37 | 41 | 1904 | 1935.23 | 31.23
42 | 1923.7 | 1935.23 | 11.53 | 42 | 1919.6 | 1935.23 | 15.63
43 | 1912.3 | 1935.23 | 22.93 | 43 | 1915.9 | 1935.23 | 19.33
44 | 1931.8 | 1935.23 | 3.43 | 44 | 1924.2 | 1935.23 | 11.03
45 | 1944.2 | 1935.23 | 8.97 | 45 | 1908.1 | 1935.23 | 27.13
46 | 1925.6 | 1935.23 | 9.63 | 46 | 1915.7 | 1935.23 | 19.53
47 | 1939.5 | 1935.23 | 4.27 | 47 | 1930.2 | 1935.23 | 5.03
48 | 1948.3 | 1935.23 | 13.07 | 48 | 1927.3 | 1935.23 | 7.93
49 | 1975.4 | 1935.23 | 40.17 | 49 | 1936.2 | 1935.23 | 0.97
50 | 1931.9 | 1935.23 | 3.33 | 50 | 1944.6 | 1935.23 | 9.37
51 | 1917.8 | 1935.23 | 17.43 | 51 | 1949.2 | 1935.23 | 13.97
52 | 1976.2 | 1935.23 | 40.97 | 52 | 1968.5 | 1935.23 | 33.27
53 | 1940.5 | 1935.23 | 5.27 | 53 | 1955.3 | 1935.23 | 20.07
54 | 1952.3 | 1935.23 | 17.07 | 54 | 1961.6 | 1935.23 | 26.37
55 | 1952.2 | 1935.23 | 16.97 | 55 | 1974.5 | 1935.23 | 39.27
56 | 1931.1 | 1935.23 | 4.13 | 56 | 1935 | 1935.23 | 0.23
57 | 1946.8 | 1935.23 | 11.57 | 57 | 1970.3 | 1935.23 | 35.07
58 | 1958.3 | 1935.23 | 23.07 | 58 | 1946.5 | 1935.23 | 11.27
59 | 1983.8 | 1935.23 | 48.57 | 59 | 1956.9 | 1935.23 | 21.67
60 | 1942 | 1935.23 | 6.77 | 60 | 1962.8 | 1935.23 | 27.57
61 | 1955.9 | 1935.23 | 20.67 | 61 | 1951.4 | 1935.23 | 16.17
62 | 1928.6 | 1935.23 | 6.63 | 62 | 1969.7 | 1935.23 | 34.47
63 | 1956.9 | 1935.23 | 21.67 | 63 | 1922.6 | 1935.23 | 12.63
64 | 1934.7 | 1935.23 | 0.53 | 64 | 1955.5 | 1935.23 | 20.27
65 | 1931.3 | 1935.23 | 3.93 | 65 | 1953 | 1935.23 | 17.77
66 | 1979.5 | 1935.23 | 44.27 | 66 | 1940.8 | 1935.23 | 5.57
67 | 1932.8 | 1935.23 | 2.43 | 67 | 1956.3 | 1935.23 | 21.07
68 | 1964.5 | 1935.23 | 29.27 | 68 | 1951.4 | 1935.23 | 16.17
69 | 1938.6 | 1935.23 | 3.37 | 69 | 1956.4 | 1935.23 | 21.17
70 | 1926.7 | 1935.23 | 8.53 | 70 | 1946.7 | 1935.23 | 11.47
71 | 1911.3 | 1935.23 | 23.93 | 71 | 1953.6 | 1935.23 | 18.37
72 | 1941 | 1935.23 | 5.77 | 72 | 1955.1 | 1935.23 | 19.87
73 | 1944.4 | 1935.23 | 9.17 | 73 | 1938 | 1935.23 | 2.77
74 | 1928.6 | 1935.23 | 6.63 | 74 | 1938.4 | 1935.23 | 3.17
75 | 1899.6 | 1935.23 | 35.63 | 75 | 1951.2 | 1935.23 | 15.97
76 | 1930.2 | 1935.23 | 5.03 | 76 | 1955.7 | 1935.23 | 20.47
77 | 1923.2 | 1935.23 | 12.03 | 77 | 1927 | 1935.23 | 8.23
78 | 1955.4 | 1935.23 | 20.17 | 78 | 1934 | 1935.23 | 1.23
79 | 1953.5 | 1935.23 | 18.27 | 79 | 1937.4 | 1935.23 | 2.17
80 | 1948.3 | 1935.23 | 13.07 | 80 | 1937.4 | 1935.23 | 2.17
81 | 1938.3 | 1935.23 | 3.07 | 81 | 1943.5 | 1935.23 | 8.27
82 | 1933.3 | 1935.23 | 1.93 | 82 | 1943 | 1935.23 | 7.77
83 | 1924.2 | 1935.23 | 11.03 | 83 | 1942.5 | 1935.23 | 7.27
84 | 1945.4 | 1935.23 | 10.17 | 84 | 1931.3 | 1935.23 | 3.93
85 | 1925.6 | 1935.23 | 9.63 | 85 | 1925.6 | 1935.23 | 9.63
86 | 1907.6 | 1935.23 | 27.63 | 86 | 1907.9 | 1935.23 | 27.33
87 | 1935.9 | 1935.23 | 0.67 | 87 | 1930 | 1935.23 | 5.23
88 | 1934.2 | 1935.23 | 1.03 | 88 | 1943.8 | 1935.23 | 8.57
89 | 1938.1 | 1935.23 | 2.87 | 89 | 1940.9 | 1935.23 | 5.67
90 | 1913 | 1935.23 | 22.23 | 90 | 1936.3 | 1935.23 | 1.07
91 | 1917.7 | 1935.23 | 17.53 | 91 | 1918.9 | 1935.23 | 16.33
92 | 1927.4 | 1935.23 | 7.83 | 92 | 1949.6 | 1935.23 | 14.37
93 | 1933.5 | 1935.23 | 1.73 | 93 | 1912.3 | 1935.23 | 22.93
94 | 1945.6 | 1935.23 | 10.37 | 94 | 1924.6 | 1935.23 | 10.63
95 | 1930.9 | 1935.23 | 4.33 | 95 | 1933 | 1935.23 | 2.23
96 | 1952.3 | 1935.23 | 17.07 | 96 | 1941 | 1935.23 | 5.77
97 | 1941.8 | 1935.23 | 6.57 | 97 | 1950.1 | 1935.23 | 14.87
98 | 1931.2 | 1935.23 | 4.03 | 98 | 1955.7 | 1935.23 | 20.47
99 | 1902.7 | 1935.23 | 32.53 | 99 | 1927.5 | 1935.23 | 7.73
100 | 1928.2 | 1935.23 | 7.03 | 100 | 1922.9 | 1935.23 | 12.33
101 | 1926.4 | 1935.23 | 8.83 | 101 | 1924.2 | 1935.23 | 11.03
102 | 1926.7 | 1935.23 | 8.53 | 102 | 1942.1 | 1935.23 | 6.87
103 | 1932.4 | 1935.23 | 2.83 | 103 | 1948.3 | 1935.23 | 13.07
104 | 1918.9 | 1935.23 | 16.33 | 104 | 1932.5 | 1935.23 | 2.73
105 | 1938.6 | 1935.23 | 3.37 | 105 | 1934.4 | 1935.23 | 0.83
106 | 1900.9 | 1935.23 | 34.33 | 106 | 1952.2 | 1935.23 | 16.97
107 | 1926.5 | 1935.23 | 8.73 | 107 | 1938.4 | 1935.23 | 3.17
108 | 1925.3 | 1935.23 | 9.93 | 108 | 1930.5 | 1935.23 | 4.73
109 | 1919.2 | 1935.23 | 16.03 | 109 | 1922.9 | 1935.23 | 12.33
110 | 1927.4 | 1935.23 | 7.83 | 110 | 1938.4 | 1935.23 | 3.17
111 | 1943.6 | 1935.23 | 8.37 | 111 | 1955.6 | 1935.23 | 20.37
112 | 1924.1 | 1935.23 | 11.13 | 112 | 1921.5 | 1935.23 | 13.73
113 | 1908.8 | 1935.23 | 26.43 | 113 | 1925.5 | 1935.23 | 9.73
114 | 1925 | 1935.23 | 10.23 | 114 | 1923.3 | 1935.23 | 11.93
115 | 1904.6 | 1935.23 | 30.63 | 115 | 1914.5 | 1935.23 | 20.73
116 | 1907 | 1935.23 | 28.23 | 116 | 1927.7 | 1935.23 | 7.53
117 | 1909.3 | 1935.23 | 25.93 | 117 | 1946.9 | 1935.23 | 11.67
118 | 1897 | 1935.23 | 38.23 | 118 | 1917.6 | 1935.23 | 17.63
119 | 1924.3 | 1935.23 | 10.93 | 119 | 1920.5 | 1935.23 | 14.73
120 | 1916.2 | 1935.23 | 19.03 | 120 | 1940.9 | 1935.23 | 5.67
121 | 1894.7 | 1935.23 | 40.53 | 121 | 1928.4 | 1935.23 | 6.83
122 | 1926.2 | 1935.23 | 9.03 | 122 | 1918.8 | 1935.23 | 16.43
123 | 1917.2 | 1935.23 | 18.03 | 123 | 1925.2 | 1935.23 | 10.03
124 | 1889.9 | 1935.23 | 45.33 | 124 | 1934.6 | 1935.23 | 0.63
125 | 1893.1 | 1935.23 | 42.13 | 125 | 1920.3 | 1935.23 | 14.93
126 | 1893.7 | 1935.23 | 41.53 | 126 | 1921.3 | 1935.23 | 13.93
127 | 1890.7 | 1935.23 | 44.53 | 127 | 1940.7 | 1935.23 | 5.47
128 | 1887.2 | 1935.23 | 48.03 | 128 | 1933.6 | 1935.23 | 1.63
129 | 1883.6 | 1935.23 | 51.63 | 129 | 1937.4 | 1935.23 | 2.17
130 | 1887.7 | 1935.23 | 47.53 | 130 | 1936.8 | 1935.23 | 1.57
131 | 1879.3 | 1935.23 | 55.93 | 131 | 1951.1 | 1935.23 | 15.87
132 | 1870.5 | 1935.23 | 64.73 | 132 | 1920.3 | 1935.23 | 14.93
133 | 1877.1 | 1935.23 | 58.13 | 133 | 1926.7 | 1935.23 | 8.53
134 | 1864.5 | 1935.23 | 70.73 | 134 | 1904.7 | 1935.23 | 30.53
135 | 1842.4 | 1935.23 | 92.83 | 135 | 1921.5 | 1935.23 | 13.73
136 | 1876.4 | 1935.23 | 58.83 | 136 | 1952.7 | 1935.23 | 17.47
137 | 1884 | 1935.23 | 51.23 | 137 | 1942.7 | 1935.23 | 7.47
138 | 1850.3 | 1935.23 | 84.93 | 138 | 1906.7 | 1935.23 | 28.53
139 | 1874.4 | 1935.23 | 60.83 | 139 | 1947.1 | 1935.23 | 11.87
140 | 1850.5 | 1935.23 | 84.73 | 140 | 1931.2 | 1935.23 | 4.03
141 | 1847.2 | 1935.23 | 88.03 | 141 | 1949.7 | 1935.23 | 14.47
142 | 1853.5 | 1935.23 | 81.73 | 142 | 1942.2 | 1935.23 | 6.97
143 | 1844.4 | 1935.23 | 90.83 | 143 | 1913.7 | 1935.23 | 21.53
144 | 1863.2 | 1935.23 | 72.03 | 144 | 1961.3 | 1935.23 | 26.07
145 | 1842.7 | 1935.23 | 92.53 | 145 | 1958.4 | 1935.23 | 23.17
146 | 1838.6 | 1935.23 | 96.63 | 146 | 1940.5 | 1935.23 | 5.27
147 | 1861.9 | 1935.23 | 73.33 | 147 | 1941.5 | 1935.23 | 6.27
148 | 1834.9 | 1935.23 | 100.33 | 148 | 1948.3 | 1935.23 | 13.07
149 | 1814.3 | 1935.23 | 120.93 | 149 | 1944.3 | 1935.23 | 9.07
150 | 1823.6 | 1935.23 | 111.63 | 150 | 1948.4 | 1935.23 | 13.17
151 | 1817.1 | 1935.23 | 118.13 | 151 | 1963.9 | 1935.23 | 28.67
152 | 1798.4 | 1935.23 | 136.83 | 152 | 1942.8 | 1935.23 | 7.57
153 | 1793 | 1935.23 | 142.23 | 153 | 1931.5 | 1935.23 | 3.73
154 | 1794.9 | 1935.23 | 140.33 | 154 | 1919.3 | 1935.23 | 15.93
155 | 1811.2 | 1935.23 | 124.03 | 155 | 1923.3 | 1935.23 | 11.93
156 | 1821 | 1935.23 | 114.23 | 156 | 1939.9 | 1935.23 | 4.67
157 | 1782 | 1935.23 | 153.23 | 157 | 1918.6 | 1935.23 | 16.63
158 | 1788.8 | 1935.23 | 146.43 | 158 | 1951.7 | 1935.23 | 16.47
159 | 1792.2 | 1935.23 | 143.03 | 159 | 1935.9 | 1935.23 | 0.67
160 | 1776.5 | 1935.23 | 158.73 | 160 | 1947.7 | 1935.23 | 12.47
Instructions:
1. Copy the values from BC_TRA_UNIFORMITY_HORIZONTAL.csv into cells B6:C165
2. Copy the values from BC_TRA_UNIFORMITY_VERTICAL.csv into cells G6:H165
3. Copy BC_TRA_UNIFORMITY_ROIs.png over cell S13

## BC Sag Uniformity
| Horizontal | Vertical
Fractional Uniformity | 0.9875 | 0.925
Horizontal |  |  |  | Vertical
Pixel | Profile Data | Central ROI Mean | Abs Difference | Pixel | Profile Data | Central ROI Mean | Abs Difference
1 | 1713.3 | 1937.67 | 224.37 | 1 | 1729.8 | 1937.67 | 207.87
2 | 1743.7 | 1937.67 | 193.97 | 2 | 1752.3 | 1937.67 | 185.37
3 | 1771 | 1937.67 | 166.67 | 3 | 1732.1 | 1937.67 | 205.57
4 | 1771.1 | 1937.67 | 166.57 | 4 | 1774.5 | 1937.67 | 163.17
5 | 1745.8 | 1937.67 | 191.87 | 5 | 1770.7 | 1937.67 | 166.97
6 | 1771.5 | 1937.67 | 166.17 | 6 | 1787.4 | 1937.67 | 150.27
7 | 1755.9 | 1937.67 | 181.77 | 7 | 1784.4 | 1937.67 | 153.27 | BC_SAG_UNIFORMITY_ROIs
8 | 1791.8 | 1937.67 | 145.87 | 8 | 1793.4 | 1937.67 | 144.27
9 | 1782.4 | 1937.67 | 155.27 | 9 | 1802.1 | 1937.67 | 135.57
10 | 1787 | 1937.67 | 150.67 | 10 | 1836.9 | 1937.67 | 100.77
11 | 1808.3 | 1937.67 | 129.37 | 11 | 1834.1 | 1937.67 | 103.57
12 | 1822 | 1937.67 | 115.67 | 12 | 1839.7 | 1937.67 | 97.97
13 | 1835.4 | 1937.67 | 102.27 | 13 | 1837 | 1937.67 | 100.67
14 | 1841.6 | 1937.67 | 96.07 | 14 | 1860.2 | 1937.67 | 77.47
15 | 1841.9 | 1937.67 | 95.77 | 15 | 1857.4 | 1937.67 | 80.27
16 | 1845.5 | 1937.67 | 92.17 | 16 | 1879.6 | 1937.67 | 58.07
17 | 1870.8 | 1937.67 | 66.87 | 17 | 1856.6 | 1937.67 | 81.07
18 | 1868.9 | 1937.67 | 68.77 | 18 | 1899.8 | 1937.67 | 37.87
19 | 1857.8 | 1937.67 | 79.87 | 19 | 1889.9 | 1937.67 | 47.77
20 | 1847.4 | 1937.67 | 90.27 | 20 | 1876.8 | 1937.67 | 60.87
21 | 1888.6 | 1937.67 | 49.07 | 21 | 1892 | 1937.67 | 45.67
22 | 1855.8 | 1937.67 | 81.87 | 22 | 1893.9 | 1937.67 | 43.77
23 | 1871.5 | 1937.67 | 66.17 | 23 | 1921 | 1937.67 | 16.67
24 | 1888.7 | 1937.67 | 48.97 | 24 | 1903.9 | 1937.67 | 33.77
25 | 1868 | 1937.67 | 69.67 | 25 | 1910.7 | 1937.67 | 26.97
26 | 1924.6 | 1937.67 | 13.07 | 26 | 1905.4 | 1937.67 | 32.27
27 | 1876.5 | 1937.67 | 61.17 | 27 | 1894.3 | 1937.67 | 43.37
28 | 1919.5 | 1937.67 | 18.17 | 28 | 1891.1 | 1937.67 | 46.57
29 | 1905.1 | 1937.67 | 32.57 | 29 | 1926.6 | 1937.67 | 11.07
30 | 1887.1 | 1937.67 | 50.57 | 30 | 1936 | 1937.67 | 1.67
31 | 1898.1 | 1937.67 | 39.57 | 31 | 1930.3 | 1937.67 | 7.37
32 | 1901.1 | 1937.67 | 36.57 | 32 | 1928.3 | 1937.67 | 9.37
33 | 1910.7 | 1937.67 | 26.97 | 33 | 1904.6 | 1937.67 | 33.07
34 | 1929.7 | 1937.67 | 7.97 | 34 | 1943 | 1937.67 | 5.33
35 | 1927.7 | 1937.67 | 9.97 | 35 | 1916 | 1937.67 | 21.67
36 | 1929.8 | 1937.67 | 7.87 | 36 | 1923.8 | 1937.67 | 13.87
37 | 1910.8 | 1937.67 | 26.87 | 37 | 1918.8 | 1937.67 | 18.87
38 | 1951.6 | 1937.67 | 13.93 | 38 | 1947.2 | 1937.67 | 9.53
39 | 1936.6 | 1937.67 | 1.07 | 39 | 1939.5 | 1937.67 | 1.83
40 | 1900.1 | 1937.67 | 37.57 | 40 | 1938.2 | 1937.67 | 0.53
41 | 1906.8 | 1937.67 | 30.87 | 41 | 1937.6 | 1937.67 | 0.07
42 | 1947.7 | 1937.67 | 10.03 | 42 | 1925.9 | 1937.67 | 11.77
43 | 1937.6 | 1937.67 | 0.07 | 43 | 1960.6 | 1937.67 | 22.93
44 | 1925.7 | 1937.67 | 11.97 | 44 | 1935.2 | 1937.67 | 2.47
45 | 1948.3 | 1937.67 | 10.63 | 45 | 1918.9 | 1937.67 | 18.77
46 | 1944.9 | 1937.67 | 7.23 | 46 | 1954.1 | 1937.67 | 16.43
47 | 1952 | 1937.67 | 14.33 | 47 | 1957.1 | 1937.67 | 19.43
48 | 1919.2 | 1937.67 | 18.47 | 48 | 1946.6 | 1937.67 | 8.93
49 | 1966.4 | 1937.67 | 28.73 | 49 | 1908.7 | 1937.67 | 28.97
50 | 1951.8 | 1937.67 | 14.13 | 50 | 1933 | 1937.67 | 4.67
51 | 1954.3 | 1937.67 | 16.63 | 51 | 1944.5 | 1937.67 | 6.83
52 | 1951.5 | 1937.67 | 13.83 | 52 | 1953.8 | 1937.67 | 16.13
53 | 1927.7 | 1937.67 | 9.97 | 53 | 1914.8 | 1937.67 | 22.87
54 | 1936.9 | 1937.67 | 0.77 | 54 | 1936.3 | 1937.67 | 1.37
55 | 1938 | 1937.67 | 0.33 | 55 | 1936.6 | 1937.67 | 1.07
56 | 1952.1 | 1937.67 | 14.43 | 56 | 1958.3 | 1937.67 | 20.63
57 | 1929.5 | 1937.67 | 8.17 | 57 | 1952.2 | 1937.67 | 14.53
58 | 1943.7 | 1937.67 | 6.03 | 58 | 1926.8 | 1937.67 | 10.87
59 | 1920 | 1937.67 | 17.67 | 59 | 1917.3 | 1937.67 | 20.37
60 | 1938.3 | 1937.67 | 0.63 | 60 | 1932.6 | 1937.67 | 5.07
61 | 1934.2 | 1937.67 | 3.47 | 61 | 1951 | 1937.67 | 13.33
62 | 1936.9 | 1937.67 | 0.77 | 62 | 1935.7 | 1937.67 | 1.97
63 | 1944.3 | 1937.67 | 6.63 | 63 | 1916 | 1937.67 | 21.67
64 | 1944.7 | 1937.67 | 7.03 | 64 | 1967.4 | 1937.67 | 29.73
65 | 1969.8 | 1937.67 | 32.13 | 65 | 1923.7 | 1937.67 | 13.97
66 | 1925.1 | 1937.67 | 12.57 | 66 | 1941.7 | 1937.67 | 4.03
67 | 1921.7 | 1937.67 | 15.97 | 67 | 1936.6 | 1937.67 | 1.07
68 | 1963.3 | 1937.67 | 25.63 | 68 | 1938.3 | 1937.67 | 0.63
69 | 1948.5 | 1937.67 | 10.83 | 69 | 1931.8 | 1937.67 | 5.87
70 | 1935.7 | 1937.67 | 1.97 | 70 | 1963.1 | 1937.67 | 25.43
71 | 1934.1 | 1937.67 | 3.57 | 71 | 1965.5 | 1937.67 | 27.83
72 | 1933.2 | 1937.67 | 4.47 | 72 | 1930 | 1937.67 | 7.67
73 | 1935.9 | 1937.67 | 1.77 | 73 | 1928.4 | 1937.67 | 9.27
74 | 1936.1 | 1937.67 | 1.57 | 74 | 1960.8 | 1937.67 | 23.13
75 | 1958.9 | 1937.67 | 21.23 | 75 | 1943.7 | 1937.67 | 6.03
76 | 1974.1 | 1937.67 | 36.43 | 76 | 1944.8 | 1937.67 | 7.13
77 | 1960.5 | 1937.67 | 22.83 | 77 | 1895.6 | 1937.67 | 42.07
78 | 1928 | 1937.67 | 9.67 | 78 | 1940 | 1937.67 | 2.33
79 | 1932.9 | 1937.67 | 4.77 | 79 | 1934.7 | 1937.67 | 2.97
80 | 1933.1 | 1937.67 | 4.57 | 80 | 1927 | 1937.67 | 10.67
81 | 1927.3 | 1937.67 | 10.37 | 81 | 1927.9 | 1937.67 | 9.77
82 | 1937.1 | 1937.67 | 0.57 | 82 | 1927.3 | 1937.67 | 10.37
83 | 1946.7 | 1937.67 | 9.03 | 83 | 1953.5 | 1937.67 | 15.83
84 | 1919.8 | 1937.67 | 17.87 | 84 | 1971.6 | 1937.67 | 33.93
85 | 1920.7 | 1937.67 | 16.97 | 85 | 1957.8 | 1937.67 | 20.13
86 | 1903.4 | 1937.67 | 34.27 | 86 | 1952.6 | 1937.67 | 14.93
87 | 1936.6 | 1937.67 | 1.07 | 87 | 1940.6 | 1937.67 | 2.93
88 | 1942.8 | 1937.67 | 5.13 | 88 | 1937.1 | 1937.67 | 0.57
89 | 1944.4 | 1937.67 | 6.73 | 89 | 1915.5 | 1937.67 | 22.17
90 | 1931.2 | 1937.67 | 6.47 | 90 | 1935.8 | 1937.67 | 1.87
91 | 1922.1 | 1937.67 | 15.57 | 91 | 1919.3 | 1937.67 | 18.37
92 | 1953.8 | 1937.67 | 16.13 | 92 | 1955.3 | 1937.67 | 17.63
93 | 1962.6 | 1937.67 | 24.93 | 93 | 1935.6 | 1937.67 | 2.07
94 | 1946.6 | 1937.67 | 8.93 | 94 | 1925.2 | 1937.67 | 12.47
95 | 1938.3 | 1937.67 | 0.63 | 95 | 1929.5 | 1937.67 | 8.17
96 | 1928.9 | 1937.67 | 8.77 | 96 | 1916.4 | 1937.67 | 21.27
97 | 1938.5 | 1937.67 | 0.83 | 97 | 1915.6 | 1937.67 | 22.07
98 | 1931 | 1937.67 | 6.67 | 98 | 1944.7 | 1937.67 | 7.03
99 | 1911 | 1937.67 | 26.67 | 99 | 1961.8 | 1937.67 | 24.13
100 | 1917.3 | 1937.67 | 20.37 | 100 | 1947.6 | 1937.67 | 9.93
101 | 1935.6 | 1937.67 | 2.07 | 101 | 1931 | 1937.67 | 6.67
102 | 1922.4 | 1937.67 | 15.27 | 102 | 1936 | 1937.67 | 1.67
103 | 1902.1 | 1937.67 | 35.57 | 103 | 1915.8 | 1937.67 | 21.87
104 | 1922.8 | 1937.67 | 14.87 | 104 | 1955.7 | 1937.67 | 18.03
105 | 1920.4 | 1937.67 | 17.27 | 105 | 1950 | 1937.67 | 12.33
106 | 1934.4 | 1937.67 | 3.27 | 106 | 1936.4 | 1937.67 | 1.27
107 | 1944.8 | 1937.67 | 7.13 | 107 | 1950.9 | 1937.67 | 13.23
108 | 1944.8 | 1937.67 | 7.13 | 108 | 1947.8 | 1937.67 | 10.13
109 | 1937 | 1937.67 | 0.67 | 109 | 1948.3 | 1937.67 | 10.63
110 | 1918.4 | 1937.67 | 19.27 | 110 | 1961.3 | 1937.67 | 23.63
111 | 1924.4 | 1937.67 | 13.27 | 111 | 1917 | 1937.67 | 20.67
112 | 1927 | 1937.67 | 10.67 | 112 | 1933.4 | 1937.67 | 4.27
113 | 1922.7 | 1937.67 | 14.97 | 113 | 1946.7 | 1937.67 | 9.03
114 | 1901 | 1937.67 | 36.67 | 114 | 1943.2 | 1937.67 | 5.53
115 | 1909.8 | 1937.67 | 27.87 | 115 | 1955.9 | 1937.67 | 18.23
116 | 1911.6 | 1937.67 | 26.07 | 116 | 1963 | 1937.67 | 25.33
117 | 1912.7 | 1937.67 | 24.97 | 117 | 1949.3 | 1937.67 | 11.63
118 | 1929.9 | 1937.67 | 7.77 | 118 | 1921.2 | 1937.67 | 16.47
119 | 1915.8 | 1937.67 | 21.87 | 119 | 1934.9 | 1937.67 | 2.77
120 | 1897.3 | 1937.67 | 40.37 | 120 | 1925.1 | 1937.67 | 12.57
121 | 1924.2 | 1937.67 | 13.47 | 121 | 1909.5 | 1937.67 | 28.17
122 | 1925.5 | 1937.67 | 12.17 | 122 | 1903.7 | 1937.67 | 33.97
123 | 1911.4 | 1937.67 | 26.27 | 123 | 1929 | 1937.67 | 8.67
124 | 1926.8 | 1937.67 | 10.87 | 124 | 1921.2 | 1937.67 | 16.47
125 | 1938.6 | 1937.67 | 0.93 | 125 | 1926 | 1937.67 | 11.67
126 | 1907.3 | 1937.67 | 30.37 | 126 | 1916.4 | 1937.67 | 21.27
127 | 1917.4 | 1937.67 | 20.27 | 127 | 1928.9 | 1937.67 | 8.77
128 | 1953.5 | 1937.67 | 15.83 | 128 | 1912 | 1937.67 | 25.67
129 | 1928.7 | 1937.67 | 8.97 | 129 | 1917.1 | 1937.67 | 20.57
130 | 1934.4 | 1937.67 | 3.27 | 130 | 1883.2 | 1937.67 | 54.47
131 | 1913.1 | 1937.67 | 24.57 | 131 | 1875.1 | 1937.67 | 62.57
132 | 1909.7 | 1937.67 | 27.97 | 132 | 1897.9 | 1937.67 | 39.77
133 | 1921.1 | 1937.67 | 16.57 | 133 | 1883.6 | 1937.67 | 54.07
134 | 1946.8 | 1937.67 | 9.13 | 134 | 1895.9 | 1937.67 | 41.77
135 | 1907.6 | 1937.67 | 30.07 | 135 | 1890.4 | 1937.67 | 47.27
136 | 1942 | 1937.67 | 4.33 | 136 | 1847.3 | 1937.67 | 90.37
137 | 1954.3 | 1937.67 | 16.63 | 137 | 1893.5 | 1937.67 | 44.17
138 | 1945.3 | 1937.67 | 7.63 | 138 | 1890.9 | 1937.67 | 46.77
139 | 1944 | 1937.67 | 6.33 | 139 | 1870.7 | 1937.67 | 66.97
140 | 1944 | 1937.67 | 6.33 | 140 | 1858.6 | 1937.67 | 79.07
141 | 1945.5 | 1937.67 | 7.83 | 141 | 1840.3 | 1937.67 | 97.37
142 | 1934.5 | 1937.67 | 3.17 | 142 | 1850.1 | 1937.67 | 87.57
143 | 1939 | 1937.67 | 1.33 | 143 | 1810.1 | 1937.67 | 127.57
144 | 1930.7 | 1937.67 | 6.97 | 144 | 1827.5 | 1937.67 | 110.17
145 | 1956.5 | 1937.67 | 18.83 | 145 | 1817.8 | 1937.67 | 119.87
146 | 1963.6 | 1937.67 | 25.93 | 146 | 1809.3 | 1937.67 | 128.37
147 | 1941.1 | 1937.67 | 3.43 | 147 | 1803.4 | 1937.67 | 134.27
148 | 1933.3 | 1937.67 | 4.37 | 148 | 1782.2 | 1937.67 | 155.47
149 | 1947.1 | 1937.67 | 9.43 | 149 | 1754.5 | 1937.67 | 183.17
150 | 1946.2 | 1937.67 | 8.53 | 150 | 1760.2 | 1937.67 | 177.47
151 | 1940.3 | 1937.67 | 2.63 | 151 | 1743.6 | 1937.67 | 194.07
152 | 1944.4 | 1937.67 | 6.73 | 152 | 1723 | 1937.67 | 214.67
153 | 1949.1 | 1937.67 | 11.43 | 153 | 1724.7 | 1937.67 | 212.97
154 | 1951.3 | 1937.67 | 13.63 | 154 | 1685.1 | 1937.67 | 252.57
155 | 1947.9 | 1937.67 | 10.23 | 155 | 1652.3 | 1937.67 | 285.37
156 | 1939 | 1937.67 | 1.33 | 156 | 1655 | 1937.67 | 282.67
157 | 1936.3 | 1937.67 | 1.37 | 157 | 1632.6 | 1937.67 | 305.07
158 | 1937.7 | 1937.67 | 0.03 | 158 | 1639.6 | 1937.67 | 298.07
159 | 1929 | 1937.67 | 8.67 | 159 | 1631 | 1937.67 | 306.67
160 | 1951.6 | 1937.67 | 13.93 | 160 | 1592.8 | 1937.67 | 344.87
Instructions:
1. Copy the values from BC_SAG_UNIFORMITY_HORIZONTAL.csv into cells B6:C165
2. Copy the values from BC_SAG_UNIFORMITY_VERTICAL.csv into cells G6:H165
3. Copy BC_SAG_UNIFORMITY_ROIs.png over cell S13

## BC Cor Uniformity
| Horizontal | Vertical
Fractional Uniformity | 1 | 0.89375
Horizontal |  |  |  | Vertical
Pixel | Profile Data | Central ROI Mean | Abs Difference | Pixel | Profile Data | Central ROI Mean | Abs Difference
1 | 1780.4 | 1955.015 | 174.615 | 1 | 1673.6 | 1955.015 | 281.415
2 | 1798.2 | 1955.015 | 156.815 | 2 | 1686.4 | 1955.015 | 268.615
3 | 1814.5 | 1955.015 | 140.515 | 3 | 1693.7 | 1955.015 | 261.315
4 | 1824.2 | 1955.015 | 130.815 | 4 | 1721.2 | 1955.015 | 233.815
5 | 1829.9 | 1955.015 | 125.115 | 5 | 1725 | 1955.015 | 230.015
6 | 1818.6 | 1955.015 | 136.415 | 6 | 1752.3 | 1955.015 | 202.715
7 | 1813.5 | 1955.015 | 141.515 | 7 | 1781.7 | 1955.015 | 173.315 | BC_COR_UNIFORMITY_ROIs
8 | 1848.8 | 1955.015 | 106.215 | 8 | 1784 | 1955.015 | 171.015
9 | 1841.4 | 1955.015 | 113.615 | 9 | 1779.2 | 1955.015 | 175.815
10 | 1839.6 | 1955.015 | 115.415 | 10 | 1808.4 | 1955.015 | 146.615
11 | 1847.8 | 1955.015 | 107.215 | 11 | 1797.6 | 1955.015 | 157.415
12 | 1842.7 | 1955.015 | 112.315 | 12 | 1809.6 | 1955.015 | 145.415
13 | 1868.5 | 1955.015 | 86.515 | 13 | 1797 | 1955.015 | 158.015
14 | 1873.5 | 1955.015 | 81.515 | 14 | 1833.3 | 1955.015 | 121.715
15 | 1853.2 | 1955.015 | 101.815 | 15 | 1830.8 | 1955.015 | 124.215
16 | 1855.2 | 1955.015 | 99.815 | 16 | 1840.2 | 1955.015 | 114.815
17 | 1875.1 | 1955.015 | 79.915 | 17 | 1835.6 | 1955.015 | 119.415
18 | 1878.6 | 1955.015 | 76.415 | 18 | 1854.4 | 1955.015 | 100.615
19 | 1894.8 | 1955.015 | 60.215 | 19 | 1896 | 1955.015 | 59.015
20 | 1891.5 | 1955.015 | 63.515 | 20 | 1883.6 | 1955.015 | 71.415
21 | 1865.4 | 1955.015 | 89.615 | 21 | 1902.7 | 1955.015 | 52.315
22 | 1883.5 | 1955.015 | 71.515 | 22 | 1894.2 | 1955.015 | 60.815
23 | 1886.2 | 1955.015 | 68.815 | 23 | 1892.5 | 1955.015 | 62.515
24 | 1901.3 | 1955.015 | 53.715 | 24 | 1886.6 | 1955.015 | 68.415
25 | 1905.9 | 1955.015 | 49.115 | 25 | 1916.7 | 1955.015 | 38.315
26 | 1893.4 | 1955.015 | 61.615 | 26 | 1923.1 | 1955.015 | 31.915
27 | 1877.3 | 1955.015 | 77.715 | 27 | 1909.2 | 1955.015 | 45.815
28 | 1886.9 | 1955.015 | 68.115 | 28 | 1927.7 | 1955.015 | 27.315
29 | 1893.7 | 1955.015 | 61.315 | 29 | 1933 | 1955.015 | 22.015
30 | 1894.4 | 1955.015 | 60.615 | 30 | 1911.5 | 1955.015 | 43.515
31 | 1904.3 | 1955.015 | 50.715 | 31 | 1906.3 | 1955.015 | 48.715
32 | 1927.4 | 1955.015 | 27.615 | 32 | 1952.6 | 1955.015 | 2.415
33 | 1906 | 1955.015 | 49.015 | 33 | 1922.8 | 1955.015 | 32.215
34 | 1918.7 | 1955.015 | 36.315 | 34 | 1918.1 | 1955.015 | 36.915
35 | 1919.4 | 1955.015 | 35.615 | 35 | 1919.3 | 1955.015 | 35.715
36 | 1904.3 | 1955.015 | 50.715 | 36 | 1928.6 | 1955.015 | 26.415
37 | 1940.2 | 1955.015 | 14.815 | 37 | 1943.4 | 1955.015 | 11.615
38 | 1934.8 | 1955.015 | 20.215 | 38 | 1930.3 | 1955.015 | 24.715
39 | 1936.7 | 1955.015 | 18.315 | 39 | 1925.8 | 1955.015 | 29.215
40 | 1921.7 | 1955.015 | 33.315 | 40 | 1918.6 | 1955.015 | 36.415
41 | 1948.3 | 1955.015 | 6.715 | 41 | 1943.6 | 1955.015 | 11.415
42 | 1944.3 | 1955.015 | 10.715 | 42 | 1953.4 | 1955.015 | 1.615
43 | 1945 | 1955.015 | 10.015 | 43 | 1941.8 | 1955.015 | 13.215
44 | 1948.4 | 1955.015 | 6.615 | 44 | 1980.4 | 1955.015 | 25.385
45 | 1952.8 | 1955.015 | 2.215 | 45 | 1967.7 | 1955.015 | 12.685
46 | 1932.7 | 1955.015 | 22.315 | 46 | 1945.5 | 1955.015 | 9.515
47 | 1940.3 | 1955.015 | 14.715 | 47 | 1965.2 | 1955.015 | 10.185
48 | 1945.5 | 1955.015 | 9.515 | 48 | 1947.5 | 1955.015 | 7.515
49 | 1942.8 | 1955.015 | 12.215 | 49 | 1945.8 | 1955.015 | 9.215
50 | 1951.5 | 1955.015 | 3.515 | 50 | 1964.6 | 1955.015 | 9.585
51 | 1941.3 | 1955.015 | 13.715 | 51 | 1969.1 | 1955.015 | 14.085
52 | 1932.9 | 1955.015 | 22.115 | 52 | 1965.8 | 1955.015 | 10.785
53 | 1970.7 | 1955.015 | 15.685 | 53 | 1961.2 | 1955.015 | 6.185
54 | 1958.4 | 1955.015 | 3.385 | 54 | 1955.1 | 1955.015 | 0.085
55 | 1956.4 | 1955.015 | 1.385 | 55 | 1978.7 | 1955.015 | 23.685
56 | 1949.2 | 1955.015 | 5.815 | 56 | 1951.5 | 1955.015 | 3.515
57 | 1927.1 | 1955.015 | 27.915 | 57 | 1952.7 | 1955.015 | 2.315
58 | 1965.2 | 1955.015 | 10.185 | 58 | 1960.9 | 1955.015 | 5.885
59 | 1968.7 | 1955.015 | 13.685 | 59 | 1957.8 | 1955.015 | 2.785
60 | 1954.1 | 1955.015 | 0.915 | 60 | 1972.1 | 1955.015 | 17.085
61 | 1956.7 | 1955.015 | 1.685 | 61 | 1921.7 | 1955.015 | 33.315
62 | 1953.4 | 1955.015 | 1.615 | 62 | 1940.4 | 1955.015 | 14.615
63 | 1951.5 | 1955.015 | 3.515 | 63 | 1952 | 1955.015 | 3.015
64 | 1961.4 | 1955.015 | 6.385 | 64 | 1978.5 | 1955.015 | 23.485
65 | 1953.2 | 1955.015 | 1.815 | 65 | 1926.5 | 1955.015 | 28.515
66 | 1987.7 | 1955.015 | 32.685 | 66 | 1951.1 | 1955.015 | 3.915
67 | 1943.5 | 1955.015 | 11.515 | 67 | 1975.6 | 1955.015 | 20.585
68 | 1958.1 | 1955.015 | 3.085 | 68 | 1963.7 | 1955.015 | 8.685
69 | 1949.4 | 1955.015 | 5.615 | 69 | 1961.6 | 1955.015 | 6.585
70 | 1945.3 | 1955.015 | 9.715 | 70 | 1947.6 | 1955.015 | 7.415
71 | 1940.8 | 1955.015 | 14.215 | 71 | 1955.2 | 1955.015 | 0.185
72 | 1982.4 | 1955.015 | 27.385 | 72 | 1938.9 | 1955.015 | 16.115
73 | 1950.1 | 1955.015 | 4.915 | 73 | 1960.3 | 1955.015 | 5.285
74 | 1980.1 | 1955.015 | 25.085 | 74 | 1960.6 | 1955.015 | 5.585
75 | 1944.9 | 1955.015 | 10.115 | 75 | 1947.6 | 1955.015 | 7.415
76 | 1948.9 | 1955.015 | 6.115 | 76 | 1963.2 | 1955.015 | 8.185
77 | 1963.5 | 1955.015 | 8.485 | 77 | 1933.2 | 1955.015 | 21.815
78 | 1967.4 | 1955.015 | 12.385 | 78 | 1962.7 | 1955.015 | 7.685
79 | 1946.7 | 1955.015 | 8.315 | 79 | 1964.4 | 1955.015 | 9.385
80 | 1961.4 | 1955.015 | 6.385 | 80 | 1949.5 | 1955.015 | 5.515
81 | 1932.1 | 1955.015 | 22.915 | 81 | 1949.4 | 1955.015 | 5.615
82 | 1928.9 | 1955.015 | 26.115 | 82 | 1955.6 | 1955.015 | 0.585
83 | 1955.2 | 1955.015 | 0.185 | 83 | 1953 | 1955.015 | 2.015
84 | 1945.3 | 1955.015 | 9.715 | 84 | 1925.4 | 1955.015 | 29.615
85 | 1951.8 | 1955.015 | 3.215 | 85 | 1944.8 | 1955.015 | 10.215
86 | 1957.2 | 1955.015 | 2.185 | 86 | 1969.3 | 1955.015 | 14.285
87 | 1961.4 | 1955.015 | 6.385 | 87 | 1922.5 | 1955.015 | 32.515
88 | 1945.7 | 1955.015 | 9.315 | 88 | 1979.4 | 1955.015 | 24.385
89 | 1921.3 | 1955.015 | 33.715 | 89 | 1966.4 | 1955.015 | 11.385
90 | 1945.7 | 1955.015 | 9.315 | 90 | 1965.7 | 1955.015 | 10.685
91 | 1945.8 | 1955.015 | 9.215 | 91 | 1972.2 | 1955.015 | 17.185
92 | 1928.2 | 1955.015 | 26.815 | 92 | 1941.8 | 1955.015 | 13.215
93 | 1976.9 | 1955.015 | 21.885 | 93 | 1949.6 | 1955.015 | 5.415
94 | 1960.1 | 1955.015 | 5.085 | 94 | 1943.1 | 1955.015 | 11.915
95 | 1946.1 | 1955.015 | 8.915 | 95 | 1967 | 1955.015 | 11.985
96 | 1955.2 | 1955.015 | 0.185 | 96 | 1968.4 | 1955.015 | 13.385
97 | 1960.2 | 1955.015 | 5.185 | 97 | 1960.1 | 1955.015 | 5.085
98 | 1943 | 1955.015 | 12.015 | 98 | 1965.7 | 1955.015 | 10.685
99 | 1946.4 | 1955.015 | 8.615 | 99 | 1941 | 1955.015 | 14.015
100 | 1957 | 1955.015 | 1.985 | 100 | 1940.5 | 1955.015 | 14.515
101 | 1938.8 | 1955.015 | 16.215 | 101 | 1934 | 1955.015 | 21.015
102 | 1950.4 | 1955.015 | 4.615 | 102 | 1949.1 | 1955.015 | 5.915
103 | 1945.8 | 1955.015 | 9.215 | 103 | 1933.5 | 1955.015 | 21.515
104 | 1951.6 | 1955.015 | 3.415 | 104 | 1969.7 | 1955.015 | 14.685
105 | 1932 | 1955.015 | 23.015 | 105 | 1930 | 1955.015 | 25.015
106 | 1955.9 | 1955.015 | 0.885 | 106 | 1955.9 | 1955.015 | 0.885
107 | 1960.9 | 1955.015 | 5.885 | 107 | 1987.3 | 1955.015 | 32.285
108 | 1935.8 | 1955.015 | 19.215 | 108 | 1970.7 | 1955.015 | 15.685
109 | 1920.3 | 1955.015 | 34.715 | 109 | 1971.9 | 1955.015 | 16.885
110 | 1946 | 1955.015 | 9.015 | 110 | 1951.6 | 1955.015 | 3.415
111 | 1954.3 | 1955.015 | 0.715 | 111 | 1932.3 | 1955.015 | 22.715
112 | 1936.3 | 1955.015 | 18.715 | 112 | 1945.4 | 1955.015 | 9.615
113 | 1936.8 | 1955.015 | 18.215 | 113 | 1961.8 | 1955.015 | 6.785
114 | 1948.9 | 1955.015 | 6.115 | 114 | 1947.5 | 1955.015 | 7.515
115 | 1948.3 | 1955.015 | 6.715 | 115 | 1957.7 | 1955.015 | 2.685
116 | 1924.3 | 1955.015 | 30.715 | 116 | 1962.3 | 1955.015 | 7.285
117 | 1931.4 | 1955.015 | 23.615 | 117 | 1949.6 | 1955.015 | 5.415
118 | 1914.7 | 1955.015 | 40.315 | 118 | 1912.5 | 1955.015 | 42.515
119 | 1928.9 | 1955.015 | 26.115 | 119 | 1893 | 1955.015 | 62.015
120 | 1920.7 | 1955.015 | 34.315 | 120 | 1950.1 | 1955.015 | 4.915
121 | 1913.9 | 1955.015 | 41.115 | 121 | 1950.6 | 1955.015 | 4.415
122 | 1920.4 | 1955.015 | 34.615 | 122 | 1920 | 1955.015 | 35.015
123 | 1913.3 | 1955.015 | 41.715 | 123 | 1917.1 | 1955.015 | 37.915
124 | 1922.1 | 1955.015 | 32.915 | 124 | 1929.6 | 1955.015 | 25.415
125 | 1913 | 1955.015 | 42.015 | 125 | 1903.2 | 1955.015 | 51.815
126 | 1922.8 | 1955.015 | 32.215 | 126 | 1912.7 | 1955.015 | 42.315
127 | 1931.6 | 1955.015 | 23.415 | 127 | 1929 | 1955.015 | 26.015
128 | 1896.9 | 1955.015 | 58.115 | 128 | 1916.2 | 1955.015 | 38.815
129 | 1904.4 | 1955.015 | 50.615 | 129 | 1899.2 | 1955.015 | 55.815
130 | 1911.4 | 1955.015 | 43.615 | 130 | 1890 | 1955.015 | 65.015
131 | 1915.7 | 1955.015 | 39.315 | 131 | 1908.4 | 1955.015 | 46.615
132 | 1891.6 | 1955.015 | 63.415 | 132 | 1913.2 | 1955.015 | 41.815
133 | 1899.3 | 1955.015 | 55.715 | 133 | 1910.7 | 1955.015 | 44.315
134 | 1898 | 1955.015 | 57.015 | 134 | 1886.2 | 1955.015 | 68.815
135 | 1935 | 1955.015 | 20.015 | 135 | 1900.5 | 1955.015 | 54.515
136 | 1920.9 | 1955.015 | 34.115 | 136 | 1909.8 | 1955.015 | 45.215
137 | 1900.5 | 1955.015 | 54.515 | 137 | 1878.5 | 1955.015 | 76.515
138 | 1888.3 | 1955.015 | 66.715 | 138 | 1874 | 1955.015 | 81.015
139 | 1866 | 1955.015 | 89.015 | 139 | 1832.5 | 1955.015 | 122.515
140 | 1885 | 1955.015 | 70.015 | 140 | 1841.3 | 1955.015 | 113.715
141 | 1907.5 | 1955.015 | 47.515 | 141 | 1866.7 | 1955.015 | 88.315
142 | 1899.5 | 1955.015 | 55.515 | 142 | 1844.5 | 1955.015 | 110.515
143 | 1870.1 | 1955.015 | 84.915 | 143 | 1835.6 | 1955.015 | 119.415
144 | 1864.2 | 1955.015 | 90.815 | 144 | 1810 | 1955.015 | 145.015
145 | 1879.2 | 1955.015 | 75.815 | 145 | 1832.6 | 1955.015 | 122.415
146 | 1865.1 | 1955.015 | 89.915 | 146 | 1821.4 | 1955.015 | 133.615
147 | 1858.4 | 1955.015 | 96.615 | 147 | 1810.6 | 1955.015 | 144.415
148 | 1839.1 | 1955.015 | 115.915 | 148 | 1794.9 | 1955.015 | 160.115
149 | 1843.9 | 1955.015 | 111.115 | 149 | 1762.8 | 1955.015 | 192.215
150 | 1855.8 | 1955.015 | 99.215 | 150 | 1751.6 | 1955.015 | 203.415
151 | 1843.4 | 1955.015 | 111.615 | 151 | 1733.7 | 1955.015 | 221.315
152 | 1843.3 | 1955.015 | 111.715 | 152 | 1726.9 | 1955.015 | 228.115
153 | 1812.2 | 1955.015 | 142.815 | 153 | 1729.6 | 1955.015 | 225.415
154 | 1823.8 | 1955.015 | 131.215 | 154 | 1723.8 | 1955.015 | 231.215
155 | 1825 | 1955.015 | 130.015 | 155 | 1697.7 | 1955.015 | 257.315
156 | 1814.3 | 1955.015 | 140.715 | 156 | 1685 | 1955.015 | 270.015
157 | 1826.4 | 1955.015 | 128.615 | 157 | 1680.3 | 1955.015 | 274.715
158 | 1830.1 | 1955.015 | 124.915 | 158 | 1665.8 | 1955.015 | 289.215
159 | 1815 | 1955.015 | 140.015 | 159 | 1654.1 | 1955.015 | 300.915
160 | 1824.2 | 1955.015 | 130.815 | 160 | 1640.3 | 1955.015 | 314.715
Instructions:
1. Copy the values from BC_COR_UNIFORMITY_HORIZONTAL.csv into cells B6:C165
2. Copy the values from BC_COR_UNIFORMITY_VERTICAL.csv into cells G6:H165
3. Copy BC_COR_UNIFORMITY_ROIs.png over cell S13

## HC Uniformity
Head Coil Uniformity
x-direction | y-direction | z-direction | Average | Standard Deviation
0.49062 | 0.38438 | 0.49375 | 0.45625 | 0.16752
| Transverse | Sagittal | Coronal
Horizontal | 0.31875 | 0.3875 | 0.6625
Vertical | 0.38125 | 0.675 | 0.3125
A
P
R
L
H
F
A
P
H
F
R
L
TRA
SAG
COR
X
Y
Y
Z
X
Z
Instructions:
1. Follow the instructions included in sheets "HC Tra Uniformity", "HC Sag Uniformity" and "HC Cor Uniformity"
2. Copy the Uniformity Profile chart into the report
The x-direction is left-right, the y-direction is up-down and the z-direction is along the bore:
y
z
x

## HC Tra Uniformity
| Horizontal | Vertical
Fractional Uniformity | 0.31875 | 0.38125
Horizontal |  |  |  | Vertical
Pixel | Profile Data | Central ROI Mean | Abs Difference | Pixel | Profile Data | Central ROI Mean | Abs Difference
1 | 1992.6 | 895.045 | 1097.555 | 1 | 1441.8 | 895.045 | 546.755
2 | 1957.7 | 895.045 | 1062.655 | 2 | 1429.9 | 895.045 | 534.855
3 | 1926.1 | 895.045 | 1031.055 | 3 | 1406.3 | 895.045 | 511.255
4 | 1893.1 | 895.045 | 998.055 | 4 | 1403.4 | 895.045 | 508.355
5 | 1863.9 | 895.045 | 968.855 | 5 | 1376.8 | 895.045 | 481.755
6 | 1830.3 | 895.045 | 935.255 | 6 | 1373.2 | 895.045 | 478.155
7 | 1802.8 | 895.045 | 907.755 | 7 | 1354.2 | 895.045 | 459.155 | HC_TRA_UNIFORMITY_ROIs
8 | 1774.3 | 895.045 | 879.255 | 8 | 1346.8 | 895.045 | 451.755
9 | 1746.3 | 895.045 | 851.255 | 9 | 1324.4 | 895.045 | 429.355
10 | 1715.4 | 895.045 | 820.355 | 10 | 1321.5 | 895.045 | 426.455
11 | 1691.2 | 895.045 | 796.155 | 11 | 1301.1 | 895.045 | 406.055
12 | 1664.1 | 895.045 | 769.055 | 12 | 1295.9 | 895.045 | 400.855
13 | 1639.4 | 895.045 | 744.355 | 13 | 1276.2 | 895.045 | 381.155
14 | 1609.8 | 895.045 | 714.755 | 14 | 1267.2 | 895.045 | 372.155
15 | 1590.3 | 895.045 | 695.255 | 15 | 1252.7 | 895.045 | 357.655
16 | 1565.1 | 895.045 | 670.055 | 16 | 1245.5 | 895.045 | 350.455
17 | 1545 | 895.045 | 649.955 | 17 | 1228.5 | 895.045 | 333.455
18 | 1518.5 | 895.045 | 623.455 | 18 | 1221.7 | 895.045 | 326.655
19 | 1498.1 | 895.045 | 603.055 | 19 | 1206.7 | 895.045 | 311.655
20 | 1475.2 | 895.045 | 580.155 | 20 | 1199.6 | 895.045 | 304.555
21 | 1456.9 | 895.045 | 561.855 | 21 | 1187.2 | 895.045 | 292.155
22 | 1432.3 | 895.045 | 537.255 | 22 | 1178.9 | 895.045 | 283.855
23 | 1414.6 | 895.045 | 519.555 | 23 | 1166.9 | 895.045 | 271.855
24 | 1394.7 | 895.045 | 499.655 | 24 | 1156.1 | 895.045 | 261.055
25 | 1376.5 | 895.045 | 481.455 | 25 | 1141.4 | 895.045 | 246.355
26 | 1353 | 895.045 | 457.955 | 26 | 1138.5 | 895.045 | 243.455
27 | 1337 | 895.045 | 441.955 | 27 | 1126.8 | 895.045 | 231.755
28 | 1319.2 | 895.045 | 424.155 | 28 | 1118.6 | 895.045 | 223.555
29 | 1304.4 | 895.045 | 409.355 | 29 | 1104 | 895.045 | 208.955
30 | 1286.3 | 895.045 | 391.255 | 30 | 1100.4 | 895.045 | 205.355
31 | 1271.1 | 895.045 | 376.055 | 31 | 1086.7 | 895.045 | 191.655
32 | 1251.6 | 895.045 | 356.555 | 32 | 1083.3 | 895.045 | 188.255
33 | 1240.2 | 895.045 | 345.155 | 33 | 1073.8 | 895.045 | 178.755
34 | 1222.1 | 895.045 | 327.055 | 34 | 1066.4 | 895.045 | 171.355
35 | 1208.2 | 895.045 | 313.155 | 35 | 1052.2 | 895.045 | 157.155
36 | 1193.9 | 895.045 | 298.855 | 36 | 1050.6 | 895.045 | 155.555
37 | 1179.4 | 895.045 | 284.355 | 37 | 1039.5 | 895.045 | 144.455
38 | 1161 | 895.045 | 265.955 | 38 | 1033.3 | 895.045 | 138.255
39 | 1152.6 | 895.045 | 257.555 | 39 | 1025.7 | 895.045 | 130.655
40 | 1137.5 | 895.045 | 242.455 | 40 | 1019.1 | 895.045 | 124.055
41 | 1126 | 895.045 | 230.955 | 41 | 1012.9 | 895.045 | 117.855
42 | 1113.2 | 895.045 | 218.155 | 42 | 1003.4 | 895.045 | 108.355
43 | 1103.2 | 895.045 | 208.155 | 43 | 999.3 | 895.045 | 104.255
44 | 1089.7 | 895.045 | 194.655 | 44 | 993.1 | 895.045 | 98.055
45 | 1080.2 | 895.045 | 185.155 | 45 | 985.2 | 895.045 | 90.155
46 | 1070.6 | 895.045 | 175.555 | 46 | 981.9 | 895.045 | 86.855
47 | 1057.8 | 895.045 | 162.755 | 47 | 974 | 895.045 | 78.955
48 | 1045.9 | 895.045 | 150.855 | 48 | 969.9 | 895.045 | 74.855
49 | 1037.1 | 895.045 | 142.055 | 49 | 963.5 | 895.045 | 68.455
50 | 1027.2 | 895.045 | 132.155 | 50 | 958.8 | 895.045 | 63.755
51 | 1019.4 | 895.045 | 124.355 | 51 | 953.6 | 895.045 | 58.555
52 | 1008.9 | 895.045 | 113.855 | 52 | 949.7 | 895.045 | 54.655
53 | 1004.7 | 895.045 | 109.655 | 53 | 942.2 | 895.045 | 47.155
54 | 993.2 | 895.045 | 98.155 | 54 | 938.5 | 895.045 | 43.455
55 | 988.2 | 895.045 | 93.155 | 55 | 932.1 | 895.045 | 37.055
56 | 979.9 | 895.045 | 84.855 | 56 | 930.7 | 895.045 | 35.655
57 | 970.3 | 895.045 | 75.255 | 57 | 926.5 | 895.045 | 31.455
58 | 962.9 | 895.045 | 67.855 | 58 | 922 | 895.045 | 26.955
59 | 959 | 895.045 | 63.955 | 59 | 916.3 | 895.045 | 21.255
60 | 950.5 | 895.045 | 55.455 | 60 | 915.6 | 895.045 | 20.555
61 | 944.7 | 895.045 | 49.655 | 61 | 913.1 | 895.045 | 18.055
62 | 935.9 | 895.045 | 40.855 | 62 | 907.1 | 895.045 | 12.055
63 | 932.3 | 895.045 | 37.255 | 63 | 907.3 | 895.045 | 12.255
64 | 927.1 | 895.045 | 32.055 | 64 | 899.5 | 895.045 | 4.455
65 | 921.7 | 895.045 | 26.655 | 65 | 899.7 | 895.045 | 4.655
66 | 916.8 | 895.045 | 21.755 | 66 | 897.9 | 895.045 | 2.855
67 | 916.1 | 895.045 | 21.055 | 67 | 897.7 | 895.045 | 2.655
68 | 914.4 | 895.045 | 19.355 | 68 | 894.4 | 895.045 | 0.645
69 | 909 | 895.045 | 13.955 | 69 | 894.8 | 895.045 | 0.245
70 | 906.1 | 895.045 | 11.055 | 70 | 890.1 | 895.045 | 4.945
71 | 905.7 | 895.045 | 10.655 | 71 | 893.4 | 895.045 | 1.645
72 | 899.2 | 895.045 | 4.155 | 72 | 887.8 | 895.045 | 7.245
73 | 896.9 | 895.045 | 1.855 | 73 | 887.3 | 895.045 | 7.745
74 | 892.8 | 895.045 | 2.245 | 74 | 885.3 | 895.045 | 9.745
75 | 892.7 | 895.045 | 2.345 | 75 | 886.8 | 895.045 | 8.245
76 | 890 | 895.045 | 5.045 | 76 | 885.8 | 895.045 | 9.245
77 | 890.6 | 895.045 | 4.445 | 77 | 889.7 | 895.045 | 5.345
78 | 891.1 | 895.045 | 3.945 | 78 | 883.9 | 895.045 | 11.145
79 | 885.6 | 895.045 | 9.445 | 79 | 888 | 895.045 | 7.045
80 | 884.1 | 895.045 | 10.945 | 80 | 886.3 | 895.045 | 8.745
81 | 884.3 | 895.045 | 10.745 | 81 | 889.1 | 895.045 | 5.945
82 | 887 | 895.045 | 8.045 | 82 | 887.8 | 895.045 | 7.245
83 | 892.3 | 895.045 | 2.745 | 83 | 888.2 | 895.045 | 6.845
84 | 887.4 | 895.045 | 7.645 | 84 | 889 | 895.045 | 6.045
85 | 888.7 | 895.045 | 6.345 | 85 | 893.3 | 895.045 | 1.745
86 | 891.4 | 895.045 | 3.645 | 86 | 895.5 | 895.045 | 0.455
87 | 895.9 | 895.045 | 0.855 | 87 | 897.7 | 895.045 | 2.655
88 | 895.7 | 895.045 | 0.655 | 88 | 898.9 | 895.045 | 3.855
89 | 899 | 895.045 | 3.955 | 89 | 903.5 | 895.045 | 8.455
90 | 899.5 | 895.045 | 4.455 | 90 | 904.1 | 895.045 | 9.055
91 | 902.7 | 895.045 | 7.655 | 91 | 907.4 | 895.045 | 12.355
92 | 906.4 | 895.045 | 11.355 | 92 | 911.3 | 895.045 | 16.255
93 | 909.3 | 895.045 | 14.255 | 93 | 916.4 | 895.045 | 21.355
94 | 911.7 | 895.045 | 16.655 | 94 | 915.7 | 895.045 | 20.655
95 | 916.5 | 895.045 | 21.455 | 95 | 921.2 | 895.045 | 26.155
96 | 921.8 | 895.045 | 26.755 | 96 | 922.9 | 895.045 | 27.855
97 | 923.7 | 895.045 | 28.655 | 97 | 929 | 895.045 | 33.955
98 | 932.5 | 895.045 | 37.455 | 98 | 931.2 | 895.045 | 36.155
99 | 938.5 | 895.045 | 43.455 | 99 | 940 | 895.045 | 44.955
100 | 938.3 | 895.045 | 43.255 | 100 | 941.8 | 895.045 | 46.755
101 | 948.4 | 895.045 | 53.355 | 101 | 950.2 | 895.045 | 55.155
102 | 953 | 895.045 | 57.955 | 102 | 951.9 | 895.045 | 56.855
103 | 960 | 895.045 | 64.955 | 103 | 963.2 | 895.045 | 68.155
104 | 967.3 | 895.045 | 72.255 | 104 | 964.5 | 895.045 | 69.455
105 | 974.8 | 895.045 | 79.755 | 105 | 975.9 | 895.045 | 80.855
106 | 983.3 | 895.045 | 88.255 | 106 | 977.6 | 895.045 | 82.555
107 | 993.2 | 895.045 | 98.155 | 107 | 991.9 | 895.045 | 96.855
108 | 999.1 | 895.045 | 104.055 | 108 | 994.6 | 895.045 | 99.555
109 | 1008.4 | 895.045 | 113.355 | 109 | 1001.1 | 895.045 | 106.055
110 | 1016.3 | 895.045 | 121.255 | 110 | 1008.3 | 895.045 | 113.255
111 | 1031.4 | 895.045 | 136.355 | 111 | 1020.6 | 895.045 | 125.555
112 | 1039.2 | 895.045 | 144.155 | 112 | 1027 | 895.045 | 131.955
113 | 1049.9 | 895.045 | 154.855 | 113 | 1038.9 | 895.045 | 143.855
114 | 1058.8 | 895.045 | 163.755 | 114 | 1040.9 | 895.045 | 145.855
115 | 1070 | 895.045 | 174.955 | 115 | 1056.9 | 895.045 | 161.855
116 | 1080.3 | 895.045 | 185.255 | 116 | 1066.6 | 895.045 | 171.555
117 | 1097.5 | 895.045 | 202.455 | 117 | 1081.5 | 895.045 | 186.455
118 | 1110.3 | 895.045 | 215.255 | 118 | 1085.1 | 895.045 | 190.055
119 | 1120.9 | 895.045 | 225.855 | 119 | 1101 | 895.045 | 205.955
120 | 1134.5 | 895.045 | 239.455 | 120 | 1108.4 | 895.045 | 213.355
121 | 1147.4 | 895.045 | 252.355 | 121 | 1123.7 | 895.045 | 228.655
122 | 1158.2 | 895.045 | 263.155 | 122 | 1131.9 | 895.045 | 236.855
123 | 1177.5 | 895.045 | 282.455 | 123 | 1148.6 | 895.045 | 253.555
124 | 1192.4 | 895.045 | 297.355 | 124 | 1157.9 | 895.045 | 262.855
125 | 1206.5 | 895.045 | 311.455 | 125 | 1176.2 | 895.045 | 281.155
126 | 1221.3 | 895.045 | 326.255 | 126 | 1184.8 | 895.045 | 289.755
127 | 1240.2 | 895.045 | 345.155 | 127 | 1200.2 | 895.045 | 305.155
128 | 1251 | 895.045 | 355.955 | 128 | 1214.8 | 895.045 | 319.755
129 | 1273.4 | 895.045 | 378.355 | 129 | 1229.9 | 895.045 | 334.855
130 | 1288.2 | 895.045 | 393.155 | 130 | 1242.8 | 895.045 | 347.755
131 | 1308.9 | 895.045 | 413.855 | 131 | 1265.1 | 895.045 | 370.055
132 | 1327 | 895.045 | 431.955 | 132 | 1277 | 895.045 | 381.955
133 | 1348.7 | 895.045 | 453.655 | 133 | 1298.3 | 895.045 | 403.255
134 | 1365.5 | 895.045 | 470.455 | 134 | 1314.9 | 895.045 | 419.855
135 | 1385.2 | 895.045 | 490.155 | 135 | 1336 | 895.045 | 440.955
136 | 1404.2 | 895.045 | 509.155 | 136 | 1350.8 | 895.045 | 455.755
137 | 1428.1 | 895.045 | 533.055 | 137 | 1373.5 | 895.045 | 478.455
138 | 1448.7 | 895.045 | 553.655 | 138 | 1391.5 | 895.045 | 496.455
139 | 1472.3 | 895.045 | 577.255 | 139 | 1415 | 895.045 | 519.955
140 | 1491.8 | 895.045 | 596.755 | 140 | 1432.7 | 895.045 | 537.655
141 | 1519.9 | 895.045 | 624.855 | 141 | 1458.2 | 895.045 | 563.155
142 | 1541.5 | 895.045 | 646.455 | 142 | 1480.6 | 895.045 | 585.555
143 | 1568.7 | 895.045 | 673.655 | 143 | 1504.7 | 895.045 | 609.655
144 | 1587.5 | 895.045 | 692.455 | 144 | 1525 | 895.045 | 629.955
145 | 1616.9 | 895.045 | 721.855 | 145 | 1553.2 | 895.045 | 658.155
146 | 1643.6 | 895.045 | 748.555 | 146 | 1575.5 | 895.045 | 680.455
147 | 1670.5 | 895.045 | 775.455 | 147 | 1598.6 | 895.045 | 703.555
148 | 1695.5 | 895.045 | 800.455 | 148 | 1624.6 | 895.045 | 729.555
149 | 1728.5 | 895.045 | 833.455 | 149 | 1651.5 | 895.045 | 756.455
150 | 1754.8 | 895.045 | 859.755 | 150 | 1678.8 | 895.045 | 783.755
151 | 1786.5 | 895.045 | 891.455 | 151 | 1709.6 | 895.045 | 814.555
152 | 1817 | 895.045 | 921.955 | 152 | 1739.5 | 895.045 | 844.455
153 | 1848.5 | 895.045 | 953.455 | 153 | 1773 | 895.045 | 877.955
154 | 1880.6 | 895.045 | 985.555 | 154 | 1802.8 | 895.045 | 907.755
155 | 1916.2 | 895.045 | 1021.155 | 155 | 1829.7 | 895.045 | 934.655
156 | 1948.9 | 895.045 | 1053.855 | 156 | 1866.3 | 895.045 | 971.255
157 | 1979.6 | 895.045 | 1084.555 | 157 | 1900.5 | 895.045 | 1005.455
158 | 2013.9 | 895.045 | 1118.855 | 158 | 1938.6 | 895.045 | 1043.555
159 | 2050.3 | 895.045 | 1155.255 | 159 | 1971.1 | 895.045 | 1076.055
160 | 2090.4 | 895.045 | 1195.355 | 160 | 2010 | 895.045 | 1114.955
Instructions:
1. Copy the values from HC_TRA_UNIFORMITY_HORIZONTAL.csv into cells B6:C165
2. Copy the values from HC_TRA_UNIFORMITY_VERTICAL.csv into cells G6:H165
3. Copy HC_TRA_UNIFORMITY_ROIs.png over cell S13

## HC Sag Uniformity
| Horizontal | Vertical
Fractional Uniformity | 0.3875 | 0.675
Horizontal |  |  |  | Vertical
Pixel | Profile Data | Central ROI Mean | Abs Difference | Pixel | Profile Data | Central ROI Mean | Abs Difference
1 | 1391.7 | 886.612 | 505.088 | 1 | 1375.5 | 886.612 | 488.888
2 | 1368.2 | 886.612 | 481.588 | 2 | 1366.8 | 886.612 | 480.188
3 | 1361.6 | 886.612 | 474.988 | 3 | 1363.3 | 886.612 | 476.688
4 | 1339.1 | 886.612 | 452.488 | 4 | 1353.2 | 886.612 | 466.588
5 | 1334 | 886.612 | 447.388 | 5 | 1347.9 | 886.612 | 461.288
6 | 1310 | 886.612 | 423.388 | 6 | 1336 | 886.612 | 449.388
7 | 1307.8 | 886.612 | 421.188 | 7 | 1332.9 | 886.612 | 446.288 | HC_SAG_UNIFORMITY_ROIs
8 | 1286 | 886.612 | 399.388 | 8 | 1320.9 | 886.612 | 434.288
9 | 1280.4 | 886.612 | 393.788 | 9 | 1318.1 | 886.612 | 431.488
10 | 1260.6 | 886.612 | 373.988 | 10 | 1305.2 | 886.612 | 418.588
11 | 1256.9 | 886.612 | 370.288 | 11 | 1302.7 | 886.612 | 416.088
12 | 1238.3 | 886.612 | 351.688 | 12 | 1287.1 | 886.612 | 400.488
13 | 1234.5 | 886.612 | 347.888 | 13 | 1282 | 886.612 | 395.388
14 | 1216.9 | 886.612 | 330.288 | 14 | 1267.5 | 886.612 | 380.888
15 | 1211.3 | 886.612 | 324.688 | 15 | 1263.7 | 886.612 | 377.088
16 | 1195.7 | 886.612 | 309.088 | 16 | 1250.5 | 886.612 | 363.888
17 | 1187.9 | 886.612 | 301.288 | 17 | 1244.8 | 886.612 | 358.188
18 | 1174.4 | 886.612 | 287.788 | 18 | 1232.5 | 886.612 | 345.888
19 | 1169.6 | 886.612 | 282.988 | 19 | 1229.2 | 886.612 | 342.588
20 | 1151.3 | 886.612 | 264.688 | 20 | 1215.2 | 886.612 | 328.588
21 | 1149 | 886.612 | 262.388 | 21 | 1210.5 | 886.612 | 323.888
22 | 1134.3 | 886.612 | 247.688 | 22 | 1199.3 | 886.612 | 312.688
23 | 1124.6 | 886.612 | 237.988 | 23 | 1190.9 | 886.612 | 304.288
24 | 1112.2 | 886.612 | 225.588 | 24 | 1181 | 886.612 | 294.388
25 | 1108.8 | 886.612 | 222.188 | 25 | 1176.1 | 886.612 | 289.488
26 | 1089.8 | 886.612 | 203.188 | 26 | 1161.8 | 886.612 | 275.188
27 | 1089 | 886.612 | 202.388 | 27 | 1156.1 | 886.612 | 269.488
28 | 1078.8 | 886.612 | 192.188 | 28 | 1144.1 | 886.612 | 257.488
29 | 1070.9 | 886.612 | 184.288 | 29 | 1138.7 | 886.612 | 252.088
30 | 1059.9 | 886.612 | 173.288 | 30 | 1127.4 | 886.612 | 240.788
31 | 1052.1 | 886.612 | 165.488 | 31 | 1124.3 | 886.612 | 237.688
32 | 1041.7 | 886.612 | 155.088 | 32 | 1111.2 | 886.612 | 224.588
33 | 1034.6 | 886.612 | 147.988 | 33 | 1104.8 | 886.612 | 218.188
34 | 1029.2 | 886.612 | 142.588 | 34 | 1094.4 | 886.612 | 207.788
35 | 1022 | 886.612 | 135.388 | 35 | 1089.8 | 886.612 | 203.188
36 | 1011.4 | 886.612 | 124.788 | 36 | 1081.6 | 886.612 | 194.988
37 | 1007 | 886.612 | 120.388 | 37 | 1072.4 | 886.612 | 185.788
38 | 1000.4 | 886.612 | 113.788 | 38 | 1064 | 886.612 | 177.388
39 | 991 | 886.612 | 104.388 | 39 | 1061.2 | 886.612 | 174.588
40 | 987.8 | 886.612 | 101.188 | 40 | 1051.7 | 886.612 | 165.088
41 | 981.8 | 886.612 | 95.188 | 41 | 1048.1 | 886.612 | 161.488
42 | 973.6 | 886.612 | 86.988 | 42 | 1038.2 | 886.612 | 151.588
43 | 970.3 | 886.612 | 83.688 | 43 | 1031.2 | 886.612 | 144.588
44 | 960 | 886.612 | 73.388 | 44 | 1026.7 | 886.612 | 140.088
45 | 954.2 | 886.612 | 67.588 | 45 | 1021.3 | 886.612 | 134.688
46 | 951.4 | 886.612 | 64.788 | 46 | 1013.9 | 886.612 | 127.288
47 | 947 | 886.612 | 60.388 | 47 | 1009.7 | 886.612 | 123.088
48 | 943 | 886.612 | 56.388 | 48 | 998.7 | 886.612 | 112.088
49 | 934.9 | 886.612 | 48.288 | 49 | 997.9 | 886.612 | 111.288
50 | 933.9 | 886.612 | 47.288 | 50 | 986.2 | 886.612 | 99.588
51 | 928.2 | 886.612 | 41.588 | 51 | 985.2 | 886.612 | 98.588
52 | 927.5 | 886.612 | 40.888 | 52 | 979.3 | 886.612 | 92.688
53 | 918.5 | 886.612 | 31.888 | 53 | 972.6 | 886.612 | 85.988
54 | 917.7 | 886.612 | 31.088 | 54 | 965.1 | 886.612 | 78.488
55 | 910.2 | 886.612 | 23.588 | 55 | 964.3 | 886.612 | 77.688
56 | 908.2 | 886.612 | 21.588 | 56 | 956.7 | 886.612 | 70.088
57 | 907.1 | 886.612 | 20.488 | 57 | 956.5 | 886.612 | 69.888
58 | 901 | 886.612 | 14.388 | 58 | 948.5 | 886.612 | 61.888
59 | 898.8 | 886.612 | 12.188 | 59 | 945.8 | 886.612 | 59.188
60 | 897.8 | 886.612 | 11.188 | 60 | 939.9 | 886.612 | 53.288
61 | 893.6 | 886.612 | 6.988 | 61 | 935.8 | 886.612 | 49.188
62 | 891.8 | 886.612 | 5.188 | 62 | 934.2 | 886.612 | 47.588
63 | 889.9 | 886.612 | 3.288 | 63 | 933.2 | 886.612 | 46.588
64 | 888.4 | 886.612 | 1.788 | 64 | 926.6 | 886.612 | 39.988
65 | 884.3 | 886.612 | 2.312 | 65 | 922.6 | 886.612 | 35.988
66 | 884.6 | 886.612 | 2.012 | 66 | 918.8 | 886.612 | 32.188
67 | 882 | 886.612 | 4.612 | 67 | 917.8 | 886.612 | 31.188
68 | 879.7 | 886.612 | 6.912 | 68 | 911.7 | 886.612 | 25.088
69 | 876.6 | 886.612 | 10.012 | 69 | 910.5 | 886.612 | 23.888
70 | 878.4 | 886.612 | 8.212 | 70 | 905.8 | 886.612 | 19.188
71 | 877 | 886.612 | 9.612 | 71 | 903.8 | 886.612 | 17.188
72 | 879.5 | 886.612 | 7.112 | 72 | 902.6 | 886.612 | 15.988
73 | 873.9 | 886.612 | 12.712 | 73 | 900.7 | 886.612 | 14.088
74 | 880.6 | 886.612 | 6.012 | 74 | 895.1 | 886.612 | 8.488
75 | 876.1 | 886.612 | 10.512 | 75 | 892.8 | 886.612 | 6.188
76 | 879.4 | 886.612 | 7.212 | 76 | 893.5 | 886.612 | 6.888
77 | 876.4 | 886.612 | 10.212 | 77 | 889.9 | 886.612 | 3.288
78 | 878.5 | 886.612 | 8.112 | 78 | 886.3 | 886.612 | 0.312
79 | 877.3 | 886.612 | 9.312 | 79 | 886.7 | 886.612 | 0.088
80 | 882.4 | 886.612 | 4.212 | 80 | 882.8 | 886.612 | 3.812
81 | 881.3 | 886.612 | 5.312 | 81 | 882.3 | 886.612 | 4.312
82 | 887.1 | 886.612 | 0.488 | 82 | 879.9 | 886.612 | 6.712
83 | 886.1 | 886.612 | 0.512 | 83 | 876.5 | 886.612 | 10.112
84 | 888.9 | 886.612 | 2.288 | 84 | 875.7 | 886.612 | 10.912
85 | 890.5 | 886.612 | 3.888 | 85 | 874.3 | 886.612 | 12.312
86 | 894.1 | 886.612 | 7.488 | 86 | 873 | 886.612 | 13.612
87 | 895.4 | 886.612 | 8.788 | 87 | 872.6 | 886.612 | 14.012
88 | 898.2 | 886.612 | 11.588 | 88 | 871.6 | 886.612 | 15.012
89 | 900.5 | 886.612 | 13.888 | 89 | 872.5 | 886.612 | 14.112
90 | 905.6 | 886.612 | 18.988 | 90 | 869.2 | 886.612 | 17.412
91 | 907.3 | 886.612 | 20.688 | 91 | 867.1 | 886.612 | 19.512
92 | 917.1 | 886.612 | 30.488 | 92 | 865 | 886.612 | 21.612
93 | 915.4 | 886.612 | 28.788 | 93 | 865.9 | 886.612 | 20.712
94 | 922.9 | 886.612 | 36.288 | 94 | 863.2 | 886.612 | 23.412
95 | 924 | 886.612 | 37.388 | 95 | 865.6 | 886.612 | 21.012
96 | 931.2 | 886.612 | 44.588 | 96 | 866 | 886.612 | 20.612
97 | 933 | 886.612 | 46.388 | 97 | 862.4 | 886.612 | 24.212
98 | 942.2 | 886.612 | 55.588 | 98 | 862.3 | 886.612 | 24.312
99 | 946.6 | 886.612 | 59.988 | 99 | 862.2 | 886.612 | 24.412
100 | 957.3 | 886.612 | 70.688 | 100 | 863 | 886.612 | 23.612
101 | 958.2 | 886.612 | 71.588 | 101 | 861.6 | 886.612 | 25.012
102 | 967 | 886.612 | 80.388 | 102 | 863.7 | 886.612 | 22.912
103 | 971.2 | 886.612 | 84.588 | 103 | 865.4 | 886.612 | 21.212
104 | 981 | 886.612 | 94.388 | 104 | 864.7 | 886.612 | 21.912
105 | 987.2 | 886.612 | 100.588 | 105 | 860.4 | 886.612 | 26.212
106 | 998.3 | 886.612 | 111.688 | 106 | 865.2 | 886.612 | 21.412
107 | 1001.2 | 886.612 | 114.588 | 107 | 864.8 | 886.612 | 21.812
108 | 1015.7 | 886.612 | 129.088 | 108 | 863.9 | 886.612 | 22.712
109 | 1020.9 | 886.612 | 134.288 | 109 | 862.9 | 886.612 | 23.712
110 | 1033.2 | 886.612 | 146.588 | 110 | 866.8 | 886.612 | 19.812
111 | 1041.3 | 886.612 | 154.688 | 111 | 865 | 886.612 | 21.612
112 | 1051.4 | 886.612 | 164.788 | 112 | 863.6 | 886.612 | 23.012
113 | 1056.9 | 886.612 | 170.288 | 113 | 866 | 886.612 | 20.612
114 | 1072.4 | 886.612 | 185.788 | 114 | 864.6 | 886.612 | 22.012
115 | 1083.3 | 886.612 | 196.688 | 115 | 864.8 | 886.612 | 21.812
116 | 1094.6 | 886.612 | 207.988 | 116 | 865.3 | 886.612 | 21.312
117 | 1103.7 | 886.612 | 217.088 | 117 | 864.7 | 886.612 | 21.912
118 | 1118.9 | 886.612 | 232.288 | 118 | 864.8 | 886.612 | 21.812
119 | 1129.7 | 886.612 | 243.088 | 119 | 864.4 | 886.612 | 22.212
120 | 1142.5 | 886.612 | 255.888 | 120 | 866.4 | 886.612 | 20.212
121 | 1154.8 | 886.612 | 268.188 | 121 | 867.8 | 886.612 | 18.812
122 | 1171 | 886.612 | 284.388 | 122 | 870.1 | 886.612 | 16.512
123 | 1179.8 | 886.612 | 293.188 | 123 | 871 | 886.612 | 15.612
124 | 1197.1 | 886.612 | 310.488 | 124 | 870.8 | 886.612 | 15.812
125 | 1210.4 | 886.612 | 323.788 | 125 | 870.2 | 886.612 | 16.412
126 | 1228.6 | 886.612 | 341.988 | 126 | 869.9 | 886.612 | 16.712
127 | 1241.4 | 886.612 | 354.788 | 127 | 869.9 | 886.612 | 16.712
128 | 1259.4 | 886.612 | 372.788 | 128 | 870.6 | 886.612 | 16.012
129 | 1274.1 | 886.612 | 387.488 | 129 | 870.2 | 886.612 | 16.412
130 | 1295.1 | 886.612 | 408.488 | 130 | 870.9 | 886.612 | 15.712
131 | 1312.7 | 886.612 | 426.088 | 131 | 871.8 | 886.612 | 14.812
132 | 1331.9 | 886.612 | 445.288 | 132 | 872 | 886.612 | 14.612
133 | 1350.7 | 886.612 | 464.088 | 133 | 870.8 | 886.612 | 15.812
134 | 1374 | 886.612 | 487.388 | 134 | 868.9 | 886.612 | 17.712
135 | 1389.3 | 886.612 | 502.688 | 135 | 869 | 886.612 | 17.612
136 | 1412.3 | 886.612 | 525.688 | 136 | 871.1 | 886.612 | 15.512
137 | 1433.6 | 886.612 | 546.988 | 137 | 870.8 | 886.612 | 15.812
138 | 1457 | 886.612 | 570.388 | 138 | 870.2 | 886.612 | 16.412
139 | 1477 | 886.612 | 590.388 | 139 | 868.5 | 886.612 | 18.112
140 | 1502.3 | 886.612 | 615.688 | 140 | 866.2 | 886.612 | 20.412
141 | 1524 | 886.612 | 637.388 | 141 | 866.1 | 886.612 | 20.512
142 | 1549.1 | 886.612 | 662.488 | 142 | 866.1 | 886.612 | 20.512
143 | 1575.1 | 886.612 | 688.488 | 143 | 865.1 | 886.612 | 21.512
144 | 1599.9 | 886.612 | 713.288 | 144 | 862.2 | 886.612 | 24.412
145 | 1628.1 | 886.612 | 741.488 | 145 | 861 | 886.612 | 25.612
146 | 1656 | 886.612 | 769.388 | 146 | 860.1 | 886.612 | 26.512
147 | 1682.2 | 886.612 | 795.588 | 147 | 855.6 | 886.612 | 31.012
148 | 1711.8 | 886.612 | 825.188 | 148 | 852.9 | 886.612 | 33.712
149 | 1746.2 | 886.612 | 859.588 | 149 | 853.7 | 886.612 | 32.912
150 | 1771.2 | 886.612 | 884.588 | 150 | 851.2 | 886.612 | 35.412
151 | 1805.8 | 886.612 | 919.188 | 151 | 848.9 | 886.612 | 37.712
152 | 1834.8 | 886.612 | 948.188 | 152 | 845.5 | 886.612 | 41.112
153 | 1873.2 | 886.612 | 986.588 | 153 | 844.9 | 886.612 | 41.712
154 | 1902.2 | 886.612 | 1015.588 | 154 | 840.7 | 886.612 | 45.912
155 | 1944.6 | 886.612 | 1057.988 | 155 | 834.4 | 886.612 | 52.212
156 | 1973.6 | 886.612 | 1086.988 | 156 | 828.7 | 886.612 | 57.912
157 | 2023.3 | 886.612 | 1136.688 | 157 | 828.4 | 886.612 | 58.212
158 | 2048.7 | 886.612 | 1162.088 | 158 | 821.2 | 886.612 | 65.412
159 | 2097 | 886.612 | 1210.388 | 159 | 817.3 | 886.612 | 69.312
160 | 2125.5 | 886.612 | 1238.888 | 160 | 811.2 | 886.612 | 75.412
Instructions:
1. Copy the values from HC_SAG_UNIFORMITY_HORIZONTAL.csv into cells B6:C165
2. Copy the values from HC_SAG_UNIFORMITY_VERTICAL.csv into cells G6:H165
3. Copy HC_SAG_UNIFORMITY_ROIs.png over cell S13

## HC Cor Uniformity
| Horizontal | Vertical
Fractional Uniformity | 0.6625 | 0.3125
Horizontal |  |  |  | Vertical
Pixel | Profile Data | Central ROI Mean | Abs Difference | Pixel | Profile Data | Central ROI Mean | Abs Difference
1 | 1400.6 | 894.795 | 505.805 | 1 | 2000 | 894.795 | 1105.205
2 | 1395.3 | 894.795 | 500.505 | 2 | 1968.7 | 894.795 | 1073.905
3 | 1389.5 | 894.795 | 494.705 | 3 | 1936.9 | 894.795 | 1042.105
4 | 1381.5 | 894.795 | 486.705 | 4 | 1907.2 | 894.795 | 1012.405
5 | 1375.1 | 894.795 | 480.305 | 5 | 1873 | 894.795 | 978.205
6 | 1367.2 | 894.795 | 472.405 | 6 | 1842.3 | 894.795 | 947.505
7 | 1362.9 | 894.795 | 468.105 | 7 | 1814.2 | 894.795 | 919.405 | HC_COR_UNIFORMITY_ROIs
8 | 1354.1 | 894.795 | 459.305 | 8 | 1785.4 | 894.795 | 890.605
9 | 1347.3 | 894.795 | 452.505 | 9 | 1754.7 | 894.795 | 859.905
10 | 1337.3 | 894.795 | 442.505 | 10 | 1725.1 | 894.795 | 830.305
11 | 1333.9 | 894.795 | 439.105 | 11 | 1700.3 | 894.795 | 805.505
12 | 1323.7 | 894.795 | 428.905 | 12 | 1673.9 | 894.795 | 779.105
13 | 1315.8 | 894.795 | 421.005 | 13 | 1645.6 | 894.795 | 750.805
14 | 1304.9 | 894.795 | 410.105 | 14 | 1620 | 894.795 | 725.205
15 | 1301.5 | 894.795 | 406.705 | 15 | 1595.2 | 894.795 | 700.405
16 | 1289.6 | 894.795 | 394.805 | 16 | 1572.9 | 894.795 | 678.105
17 | 1283.8 | 894.795 | 389.005 | 17 | 1544.4 | 894.795 | 649.605
18 | 1269 | 894.795 | 374.205 | 18 | 1521.9 | 894.795 | 627.105
19 | 1263 | 894.795 | 368.205 | 19 | 1503.3 | 894.795 | 608.505
20 | 1251.8 | 894.795 | 357.005 | 20 | 1478.6 | 894.795 | 583.805
21 | 1246.3 | 894.795 | 351.505 | 21 | 1456.4 | 894.795 | 561.605
22 | 1234.8 | 894.795 | 340.005 | 22 | 1435.7 | 894.795 | 540.905
23 | 1229.1 | 894.795 | 334.305 | 23 | 1414 | 894.795 | 519.205
24 | 1217.1 | 894.795 | 322.305 | 24 | 1394.2 | 894.795 | 499.405
25 | 1211.9 | 894.795 | 317.105 | 25 | 1372.8 | 894.795 | 478.005
26 | 1198.7 | 894.795 | 303.905 | 26 | 1356.1 | 894.795 | 461.305
27 | 1191.7 | 894.795 | 296.905 | 27 | 1338.5 | 894.795 | 443.705
28 | 1179.6 | 894.795 | 284.805 | 28 | 1322.4 | 894.795 | 427.605
29 | 1174.2 | 894.795 | 279.405 | 29 | 1300.2 | 894.795 | 405.405
30 | 1163.1 | 894.795 | 268.305 | 30 | 1287 | 894.795 | 392.205
31 | 1154.9 | 894.795 | 260.105 | 31 | 1270.2 | 894.795 | 375.405
32 | 1145.7 | 894.795 | 250.905 | 32 | 1254.7 | 894.795 | 359.905
33 | 1139.5 | 894.795 | 244.705 | 33 | 1238.4 | 894.795 | 343.605
34 | 1125.3 | 894.795 | 230.505 | 34 | 1221.3 | 894.795 | 326.505
35 | 1120.6 | 894.795 | 225.805 | 35 | 1209.3 | 894.795 | 314.505
36 | 1108.7 | 894.795 | 213.905 | 36 | 1191.3 | 894.795 | 296.505
37 | 1104.9 | 894.795 | 210.105 | 37 | 1176.6 | 894.795 | 281.805
38 | 1098.5 | 894.795 | 203.705 | 38 | 1163.1 | 894.795 | 268.305
39 | 1090 | 894.795 | 195.205 | 39 | 1150.1 | 894.795 | 255.305
40 | 1082.2 | 894.795 | 187.405 | 40 | 1137.6 | 894.795 | 242.805
41 | 1073.6 | 894.795 | 178.805 | 41 | 1122.4 | 894.795 | 227.605
42 | 1065.8 | 894.795 | 171.005 | 42 | 1115.6 | 894.795 | 220.805
43 | 1060.1 | 894.795 | 165.305 | 43 | 1101.4 | 894.795 | 206.605
44 | 1050.3 | 894.795 | 155.505 | 44 | 1087.5 | 894.795 | 192.705
45 | 1046.9 | 894.795 | 152.105 | 45 | 1077.7 | 894.795 | 182.905
46 | 1040.2 | 894.795 | 145.405 | 46 | 1065.6 | 894.795 | 170.805
47 | 1027.7 | 894.795 | 132.905 | 47 | 1059.5 | 894.795 | 164.705
48 | 1025.3 | 894.795 | 130.505 | 48 | 1047.9 | 894.795 | 153.105
49 | 1020.8 | 894.795 | 126.005 | 49 | 1035.9 | 894.795 | 141.105
50 | 1011.2 | 894.795 | 116.405 | 50 | 1027.7 | 894.795 | 132.905
51 | 1009.6 | 894.795 | 114.805 | 51 | 1018.3 | 894.795 | 123.505
52 | 999.6 | 894.795 | 104.805 | 52 | 1010 | 894.795 | 115.205
53 | 996.3 | 894.795 | 101.505 | 53 | 1003.6 | 894.795 | 108.805
54 | 991 | 894.795 | 96.205 | 54 | 993.5 | 894.795 | 98.705
55 | 983.1 | 894.795 | 88.305 | 55 | 985.7 | 894.795 | 90.905
56 | 978.2 | 894.795 | 83.405 | 56 | 978 | 894.795 | 83.205
57 | 974.3 | 894.795 | 79.505 | 57 | 969.4 | 894.795 | 74.605
58 | 966 | 894.795 | 71.205 | 58 | 963.1 | 894.795 | 68.305
59 | 964 | 894.795 | 69.205 | 59 | 958.2 | 894.795 | 63.405
60 | 959.2 | 894.795 | 64.405 | 60 | 948.9 | 894.795 | 54.105
61 | 952.7 | 894.795 | 57.905 | 61 | 945.2 | 894.795 | 50.405
62 | 950.9 | 894.795 | 56.105 | 62 | 938.2 | 894.795 | 43.405
63 | 946.1 | 894.795 | 51.305 | 63 | 932.3 | 894.795 | 37.505
64 | 938 | 894.795 | 43.205 | 64 | 927.8 | 894.795 | 33.005
65 | 937.6 | 894.795 | 42.805 | 65 | 923.9 | 894.795 | 29.105
66 | 931.3 | 894.795 | 36.505 | 66 | 918.4 | 894.795 | 23.605
67 | 931 | 894.795 | 36.205 | 67 | 914.5 | 894.795 | 19.705
68 | 924.2 | 894.795 | 29.405 | 68 | 913.9 | 894.795 | 19.105
69 | 923.7 | 894.795 | 28.905 | 69 | 910.6 | 894.795 | 15.805
70 | 915.1 | 894.795 | 20.305 | 70 | 906.3 | 894.795 | 11.505
71 | 917 | 894.795 | 22.205 | 71 | 900.3 | 894.795 | 5.505
72 | 914.3 | 894.795 | 19.505 | 72 | 900.9 | 894.795 | 6.105
73 | 908.9 | 894.795 | 14.105 | 73 | 899.2 | 894.795 | 4.405
74 | 908.3 | 894.795 | 13.505 | 74 | 895.7 | 894.795 | 0.905
75 | 904.7 | 894.795 | 9.905 | 75 | 895.1 | 894.795 | 0.305
76 | 902.6 | 894.795 | 7.805 | 76 | 889.4 | 894.795 | 5.395
77 | 897.9 | 894.795 | 3.105 | 77 | 888 | 894.795 | 6.795
78 | 893 | 894.795 | 1.795 | 78 | 889.6 | 894.795 | 5.195
79 | 894.1 | 894.795 | 0.695 | 79 | 888.5 | 894.795 | 6.295
80 | 889.7 | 894.795 | 5.095 | 80 | 887.3 | 894.795 | 7.495
81 | 887.7 | 894.795 | 7.095 | 81 | 890.7 | 894.795 | 4.095
82 | 882.6 | 894.795 | 12.195 | 82 | 892.4 | 894.795 | 2.395
83 | 883.8 | 894.795 | 10.995 | 83 | 890.5 | 894.795 | 4.295
84 | 882.5 | 894.795 | 12.295 | 84 | 889.5 | 894.795 | 5.295
85 | 881.8 | 894.795 | 12.995 | 85 | 889.8 | 894.795 | 4.995
86 | 878.3 | 894.795 | 16.495 | 86 | 887.1 | 894.795 | 7.695
87 | 880.2 | 894.795 | 14.595 | 87 | 892.9 | 894.795 | 1.895
88 | 878.8 | 894.795 | 15.995 | 88 | 896.3 | 894.795 | 1.505
89 | 875.5 | 894.795 | 19.295 | 89 | 899.6 | 894.795 | 4.805
90 | 872.5 | 894.795 | 22.295 | 90 | 898.2 | 894.795 | 3.405
91 | 870.7 | 894.795 | 24.095 | 91 | 905.3 | 894.795 | 10.505
92 | 871.9 | 894.795 | 22.895 | 92 | 907.1 | 894.795 | 12.305
93 | 870.3 | 894.795 | 24.495 | 93 | 914 | 894.795 | 19.205
94 | 868.7 | 894.795 | 26.095 | 94 | 917.7 | 894.795 | 22.905
95 | 865 | 894.795 | 29.795 | 95 | 919.1 | 894.795 | 24.305
96 | 867 | 894.795 | 27.795 | 96 | 922.6 | 894.795 | 27.805
97 | 865.1 | 894.795 | 29.695 | 97 | 927 | 894.795 | 32.205
98 | 865.4 | 894.795 | 29.395 | 98 | 931.9 | 894.795 | 37.105
99 | 865 | 894.795 | 29.795 | 99 | 940 | 894.795 | 45.205
100 | 861.1 | 894.795 | 33.695 | 100 | 943.4 | 894.795 | 48.605
101 | 862.6 | 894.795 | 32.195 | 101 | 949.7 | 894.795 | 54.905
102 | 861.6 | 894.795 | 33.195 | 102 | 953.4 | 894.795 | 58.605
103 | 865 | 894.795 | 29.795 | 103 | 965.5 | 894.795 | 70.705
104 | 864.9 | 894.795 | 29.895 | 104 | 972 | 894.795 | 77.205
105 | 864.3 | 894.795 | 30.495 | 105 | 980.4 | 894.795 | 85.605
106 | 860.6 | 894.795 | 34.195 | 106 | 985.2 | 894.795 | 90.405
107 | 863.2 | 894.795 | 31.595 | 107 | 992.6 | 894.795 | 97.805
108 | 862.4 | 894.795 | 32.395 | 108 | 1000.6 | 894.795 | 105.805
109 | 861.4 | 894.795 | 33.395 | 109 | 1011.9 | 894.795 | 117.105
110 | 862.5 | 894.795 | 32.295 | 110 | 1016.5 | 894.795 | 121.705
111 | 864.5 | 894.795 | 30.295 | 111 | 1028.9 | 894.795 | 134.105
112 | 861.5 | 894.795 | 33.295 | 112 | 1036.7 | 894.795 | 141.905
113 | 861.5 | 894.795 | 33.295 | 113 | 1051.7 | 894.795 | 156.905
114 | 860.5 | 894.795 | 34.295 | 114 | 1056.4 | 894.795 | 161.605
115 | 861.4 | 894.795 | 33.395 | 115 | 1071.2 | 894.795 | 176.405
116 | 863.6 | 894.795 | 31.195 | 116 | 1082.7 | 894.795 | 187.905
117 | 863.5 | 894.795 | 31.295 | 117 | 1096.9 | 894.795 | 202.105
118 | 862.8 | 894.795 | 31.995 | 118 | 1105.8 | 894.795 | 211.005
119 | 861.3 | 894.795 | 33.495 | 119 | 1122.9 | 894.795 | 228.105
120 | 863.6 | 894.795 | 31.195 | 120 | 1130.4 | 894.795 | 235.605
121 | 862.9 | 894.795 | 31.895 | 121 | 1151.6 | 894.795 | 256.805
122 | 864.9 | 894.795 | 29.895 | 122 | 1159.5 | 894.795 | 264.705
123 | 866.9 | 894.795 | 27.895 | 123 | 1178.5 | 894.795 | 283.705
124 | 867.5 | 894.795 | 27.295 | 124 | 1191.4 | 894.795 | 296.605
125 | 864.8 | 894.795 | 29.995 | 125 | 1208.4 | 894.795 | 313.605
126 | 864.6 | 894.795 | 30.195 | 126 | 1220.3 | 894.795 | 325.505
127 | 867.4 | 894.795 | 27.395 | 127 | 1239.7 | 894.795 | 344.905
128 | 863.8 | 894.795 | 30.995 | 128 | 1252 | 894.795 | 357.205
129 | 864.2 | 894.795 | 30.595 | 129 | 1275 | 894.795 | 380.205
130 | 865.2 | 894.795 | 29.595 | 130 | 1290.1 | 894.795 | 395.305
131 | 866.6 | 894.795 | 28.195 | 131 | 1311.2 | 894.795 | 416.405
132 | 869.2 | 894.795 | 25.595 | 132 | 1327.4 | 894.795 | 432.605
133 | 869.5 | 894.795 | 25.295 | 133 | 1350.4 | 894.795 | 455.605
134 | 867.2 | 894.795 | 27.595 | 134 | 1364.5 | 894.795 | 469.705
135 | 865.4 | 894.795 | 29.395 | 135 | 1388.5 | 894.795 | 493.705
136 | 866 | 894.795 | 28.795 | 136 | 1406.1 | 894.795 | 511.305
137 | 867.3 | 894.795 | 27.495 | 137 | 1431.9 | 894.795 | 537.105
138 | 864.9 | 894.795 | 29.895 | 138 | 1446.5 | 894.795 | 551.705
139 | 863.7 | 894.795 | 31.095 | 139 | 1474.8 | 894.795 | 580.005
140 | 865.5 | 894.795 | 29.295 | 140 | 1494.6 | 894.795 | 599.805
141 | 864.3 | 894.795 | 30.495 | 141 | 1521.5 | 894.795 | 626.705
142 | 862.4 | 894.795 | 32.395 | 142 | 1544.2 | 894.795 | 649.405
143 | 862.7 | 894.795 | 32.095 | 143 | 1570.7 | 894.795 | 675.905
144 | 861.1 | 894.795 | 33.695 | 144 | 1591.3 | 894.795 | 696.505
145 | 860.7 | 894.795 | 34.095 | 145 | 1622.9 | 894.795 | 728.105
146 | 860.5 | 894.795 | 34.295 | 146 | 1644.5 | 894.795 | 749.705
147 | 857.4 | 894.795 | 37.395 | 147 | 1678.3 | 894.795 | 783.505
148 | 854.5 | 894.795 | 40.295 | 148 | 1700.2 | 894.795 | 805.405
149 | 851.8 | 894.795 | 42.995 | 149 | 1734.1 | 894.795 | 839.305
150 | 852.4 | 894.795 | 42.395 | 150 | 1760.9 | 894.795 | 866.105
151 | 851.3 | 894.795 | 43.495 | 151 | 1794.3 | 894.795 | 899.505
152 | 848 | 894.795 | 46.795 | 152 | 1821.5 | 894.795 | 926.705
153 | 843.6 | 894.795 | 51.195 | 153 | 1857.1 | 894.795 | 962.305
154 | 842.7 | 894.795 | 52.095 | 154 | 1884.8 | 894.795 | 990.005
155 | 837.7 | 894.795 | 57.095 | 155 | 1920.7 | 894.795 | 1025.905
156 | 839.6 | 894.795 | 55.195 | 156 | 1952.6 | 894.795 | 1057.805
157 | 831 | 894.795 | 63.795 | 157 | 1992.8 | 894.795 | 1098.005
158 | 826.6 | 894.795 | 68.195 | 158 | 2026.4 | 894.795 | 1131.605
159 | 822.3 | 894.795 | 72.495 | 159 | 2065.4 | 894.795 | 1170.605
160 | 821.1 | 894.795 | 73.695 | 160 | 2100.2 | 894.795 | 1205.405
Instructions:
1. Copy the values from HC_COR_UNIFORMITY_HORIZONTAL.csv into cells B6:C165
2. Copy the values from HC_COR_UNIFORMITY_VERTICAL.csv into cells G6:H165
3. Copy HC_COR_UNIFORMITY_ROIs.png over cell S13

## Geometric
Geometric Linearity and Distortion
Imaging Planes
| Transverse | Sagittal | Coronal | Average | Standard Deviation
Distance (mm) | 120.5822 | 120.8373 | 120.2545 | 120.558 | 0.39298
Distortion (%) | 0.44837 | 0.16598 | 0.1557 | 0.25668
Gradient Directions
| x-direction | y-direction | z-direction | Average | Standard Deviation
Distance (mm) | 120.137 | 121.0235 | 120.5135 | 120.558 | 0.39298
Distortion (%) | 0.14356 | 0.058183 | 0.20764 | 0.13646
Distance between rods (mm) | 120
Plane: Transverse |  |  |  | Plane: Sagittal |  |  |  | Plane: Coronal
| Measured Distance (mm) | Error (mm) |  |  | Measured Distance (mm) | Error (mm) |  |  | Measured Distance (mm) | Error (mm)
Rod 1 to Rod 3 | 120.004 | 0.004 |  | Rod 1 to Rod 3 | 120.946 | 0.946 |  | Rod 1 to Rod 3 | 119.95 | -0.05
Rod 4 to Rod 6 | 120.004 | 0.004 |  | Rod 4 to Rod 6 | 121.001 | 1.001 |  | Rod 4 to Rod 6 | 120.356 | 0.356
Rod 7 to Rod 9 | 120.293 | 0.293 |  | Rod 7 to Rod 9 | 121.002 | 1.002 |  | Rod 7 to Rod 9 | 120.215 | 0.215
Rod 1 to Rod 7 | 121.12 | 1.12 |  | Rod 1 to Rod 7 | 120.67 | 0.67 |  | Rod 1 to Rod 7 | 120.37 | 0.37
Rod 2 to Rod 8 | 121.1 | 1.1 |  | Rod 2 to Rod 8 | 120.514 | 0.514 |  | Rod 2 to Rod 8 | 120.475 | 0.475
Rod 3 to Rod 9 | 120.972 | 0.972 |  | Rod 3 to Rod 9 | 120.891 | 0.891 |  | Rod 3 to Rod 9 | 120.161 | 0.161
BC_TRA_GEOMETRIC_LINEARITY_AND_DISTORTION_LINES |  |  | BC_SAG_GEOMETRIC_LINEARITY_AND_DISTORTION_LINES |  |  |  | BC_COR_GEOMETRIC_LINEARITY_AND_DISTORTION_LINES
The x-direction is left-right, the y-direction is up-down and the z-direction is along the bore:
y
z
x
Rod positions:
1
2
3
4
5
6
7
8
9
Instructions:
1. Check the distance between the rods in cell B13, it should be 120 mm
2. Copy the values from BC_TRA_GEOMETRIC_LINEARITY_DISTORTION.csv into cells B17:B22
3. Copy the values from BC_SAG_GEOMETRIC_LINEARITY_DISTORTION.csv into cells E17:E22
4. Copy the values from BC_COR_GEOMETRIC_LINEARITY_DISTORTION.csv into cells F17:F22
5. Copy BC_TRA_GEOMETRIC_LINEARITY_AND_DISTORTION_LINES.png over cell A38
6. Copy BC_SAG_GEOMETRIC_LINEARITY_AND_DISTORTION_LINES.png over cell D38
7. Copy BC_COR_GEOMETRIC_LINEARITY_AND_DISTORTION_LINES.png over cell H38

## Slice Width
Slice Width
Slice width (mm)
Nominal | Transverse | Sagittal | Coronal | Average | Standard Deviation
5 | 5.1953 | 5.1953 | 5.1953 | 5.1953 | 0
Input Parameters | Value
Frequency encode matrix size | 256
Field-of-view (mm) | 256
Fixed Parameters | Value
Angle of plates in test object (deg) | 11.3
Calculated Parameters | Value
Stretch factor | 5.0045
Voxel size in the frequency encode direction (mm) | 1
FWHM transverse profile 1 (pixels) | 26
FWHM transverse profile 2 (pixels) | 26
FWHM sagittal profile 1 (pixels) | 26
FWHM sagittal profile 2 (pixels) | 26
FWHM coronal profile 1 (pixels) | 26
FWHM coronal profile 2 (pixels) | 26
Instructions:
1. Enter the frequency encode matrix size into cell B8
2. Enter the field-of-view into cell B9
3. Follow the instructions included in sheets "Tra Slice Width", "Sag Slice Width" and " Cor Slice Width"

## Tra Slice Width
Fitting Parameters | Value
Manual offset for baseline of profile 1 | 0
Manual offset for minimum of profile 1 | 25
Manual offset for baseline of profile 2 | 0
Manual offset for minimum of profile 2 | 25
Calculated Parameters | Value
FWHM profile 1 (pixels) | 26
FWHM profile 2 (pixels) | 26 |  | BC_TRA_SLICE_WIDTH_ROIs
Pixel | Profile 1 | Profile 2
0 | 1598.2 | 1666.4
1 | 1595.2 | 1668
2 | 1620 | 1644.3
3 | 1638.3 | 1638.6
4 | 1612.9 | 1657.8
5 | 1613.5 | 1648.1
6 | 1625.8 | 1640.4
7 | 1618.2 | 1642.1
8 | 1610.8 | 1643
9 | 1612.2 | 1654
10 | 1619.2 | 1668.3
11 | 1618.3 | 1662.2
12 | 1624.2 | 1643.7
13 | 1620.1 | 1637.1
14 | 1601.2 | 1628.1
15 | 1620.2 | 1633.4
16 | 1651.1 | 1652
17 | 1638.2 | 1669.7
18 | 1614.7 | 1665.9
19 | 1614.5 | 1668.9
20 | 1645.3 | 1688.7
21 | 1625.1 | 1682.4
22 | 1631.7 | 1661.5
23 | 1605.4 | 1655.4
24 | 1607 | 1674.4
25 | 1635.7 | 1652.8
26 | 1649.9 | 1642.5
27 | 1608.8 | 1686.7
28 | 1628.4 | 1679.5
29 | 1623.4 | 1678.9
30 | 1643.9 | 1674.4
31 | 1634.3 | 1667.4
32 | 1624.8 | 1683.3
33 | 1606.3 | 1684.4
34 | 1633.9 | 1667
35 | 1605.4 | 1662.1
36 | 1607.3 | 1663.3
37 | 1636.1 | 1668.5
38 | 1639.2 | 1663.7
39 | 1595.3 | 1678.4
40 | 1566.4 | 1692.1
41 | 1570.8 | 1676.2
42 | 1524.9 | 1670.9
43 | 1500.9 | 1638.8
44 | 1478.9 | 1623.9
45 | 1435.2 | 1578.6
46 | 1406.5 | 1551.8
47 | 1375.9 | 1520.2
48 | 1361.7 | 1510.3
49 | 1332.1 | 1457.8
50 | 1284.7 | 1422.5
51 | 1288.2 | 1419.8
52 | 1293.7 | 1361.1
53 | 1263.5 | 1321.1
54 | 1266.7 | 1311.2
55 | 1234.9 | 1321.5
56 | 1255.1 | 1296.1
57 | 1272.2 | 1278.2
58 | 1258.6 | 1282.4
59 | 1255.6 | 1246.8
60 | 1260.7 | 1249.3
61 | 1276.6 | 1277.5
62 | 1257.9 | 1274.5
63 | 1275.4 | 1283.7
64 | 1277.1 | 1281.3
65 | 1296.8 | 1273.5
66 | 1294.7 | 1266.2
67 | 1313.3 | 1307.3
68 | 1339.5 | 1286.7
69 | 1364.1 | 1310.5
70 | 1391 | 1367.4
71 | 1459.1 | 1369.5
72 | 1490.5 | 1382.2
73 | 1541.5 | 1442.6
74 | 1562.1 | 1447.6
75 | 1587.8 | 1511.3
76 | 1608 | 1536.7
77 | 1610.4 | 1546.3
78 | 1618.4 | 1586.2
79 | 1613 | 1598.4
80 | 1624.1 | 1644.8
81 | 1637.2 | 1668.3
82 | 1643.6 | 1680.4
83 | 1638.6 | 1674.5
84 | 1621.5 | 1682.7
85 | 1635.9 | 1669.1
86 | 1657.6 | 1668
87 | 1650.2 | 1673.2
88 | 1648.2 | 1662.2
89 | 1640.4 | 1697.6
90 | 1671.5 | 1699.2
91 | 1661.8 | 1677.7
92 | 1638.6 | 1690.1
93 | 1655 | 1682.4
94 | 1646.9 | 1697.5
95 | 1636.2 | 1697.1
96 | 1645.4 | 1682.2
97 | 1657 | 1665.8
98 | 1661.1 | 1676
99 | 1647.2 | 1703.5
100 | 1649.9 | 1696.7
101 | 1647 | 1681.3
102 | 1653.5 | 1697.7
103 | 1622.1 | 1678.6
104 | 1647.5 | 1684.4
105 | 1659 | 1680.7
106 | 1620.3 | 1678.3
107 | 1635.6 | 1690.1
108 | 1631.1 | 1691.9
109 | 1641.8 | 1677.3
110 | 1642 | 1672.1
111 | 1642.3 | 1672.1
112 | 1645.9 | 1659.2
113 | 1641.5 | 1666.9
114 | 1651.8 | 1671
115 | 1635.4 | 1689.5
116 | 1627.3 | 1700.3
117 | 1626.5 | 1694.3
118 | 1625.2 | 1683.7
119 | 1626.4 | 1668.9
Baseline | 1632.3323 | 1671.1823
Minimum | 1259.9 | 1271.8
Half Minimum | 1446.1161 | 1471.4911
Instructions:
1. Copy the values from BC_TRA_SLICE_WIDTH_1.csv into cells B12:B131
2. Copy the values from BC_TRA_SLICE_WIDTH_2.csv into cells C12:C131
3. Check the Baseline shown in the Profile 1 chart. If necessary adjust the position, using the offset in cell B2
4. Check the Minimum shown in the Profile 1 chart . If necessary adjust the position, using the offset in cell B3
5. Check the Baseline shown in the Profile 2 chart. If necessary adjust the position, using the offset in cell B4
6. Check the Minimum shown in the Profile 2 chart . If necessary adjust the position, using the offset in cell B5
7. Copy BC_TRA_SLICE_WIDTH_ROIs.png over cell N11

## Sag Slice Width
Fitting Parameters | Value
Manual offset for baseline of profile 1 | 0
Manual offset for minimum of profile 1 | 25
Manual offset for baseline of profile 2 | 0
Manual offset for minimum of profile 2 | 25
Calculated Parameters | Value
FWHM profile 1 (pixels) | 26
FWHM profile 2 (pixels) | 26 |  | BC_SAG_SLICE_WIDTH_ROIs
Pixel | Profile 1 | Profile 2
0 | 1598.2 | 1666.4
1 | 1595.2 | 1668
2 | 1620 | 1644.3
3 | 1638.3 | 1638.6
4 | 1612.9 | 1657.8
5 | 1613.5 | 1648.1
6 | 1625.8 | 1640.4
7 | 1618.2 | 1642.1
8 | 1610.8 | 1643
9 | 1612.2 | 1654
10 | 1619.2 | 1668.3
11 | 1618.3 | 1662.2
12 | 1624.2 | 1643.7
13 | 1620.1 | 1637.1
14 | 1601.2 | 1628.1
15 | 1620.2 | 1633.4
16 | 1651.1 | 1652
17 | 1638.2 | 1669.7
18 | 1614.7 | 1665.9
19 | 1614.5 | 1668.9
20 | 1645.3 | 1688.7
21 | 1625.1 | 1682.4
22 | 1631.7 | 1661.5
23 | 1605.4 | 1655.4
24 | 1607 | 1674.4
25 | 1635.7 | 1652.8
26 | 1649.9 | 1642.5
27 | 1608.8 | 1686.7
28 | 1628.4 | 1679.5
29 | 1623.4 | 1678.9
30 | 1643.9 | 1674.4
31 | 1634.3 | 1667.4
32 | 1624.8 | 1683.3
33 | 1606.3 | 1684.4
34 | 1633.9 | 1667
35 | 1605.4 | 1662.1
36 | 1607.3 | 1663.3
37 | 1636.1 | 1668.5
38 | 1639.2 | 1663.7
39 | 1595.3 | 1678.4
40 | 1566.4 | 1692.1
41 | 1570.8 | 1676.2
42 | 1524.9 | 1670.9
43 | 1500.9 | 1638.8
44 | 1478.9 | 1623.9
45 | 1435.2 | 1578.6
46 | 1406.5 | 1551.8
47 | 1375.9 | 1520.2
48 | 1361.7 | 1510.3
49 | 1332.1 | 1457.8
50 | 1284.7 | 1422.5
51 | 1288.2 | 1419.8
52 | 1293.7 | 1361.1
53 | 1263.5 | 1321.1
54 | 1266.7 | 1311.2
55 | 1234.9 | 1321.5
56 | 1255.1 | 1296.1
57 | 1272.2 | 1278.2
58 | 1258.6 | 1282.4
59 | 1255.6 | 1246.8
60 | 1260.7 | 1249.3
61 | 1276.6 | 1277.5
62 | 1257.9 | 1274.5
63 | 1275.4 | 1283.7
64 | 1277.1 | 1281.3
65 | 1296.8 | 1273.5
66 | 1294.7 | 1266.2
67 | 1313.3 | 1307.3
68 | 1339.5 | 1286.7
69 | 1364.1 | 1310.5
70 | 1391 | 1367.4
71 | 1459.1 | 1369.5
72 | 1490.5 | 1382.2
73 | 1541.5 | 1442.6
74 | 1562.1 | 1447.6
75 | 1587.8 | 1511.3
76 | 1608 | 1536.7
77 | 1610.4 | 1546.3
78 | 1618.4 | 1586.2
79 | 1613 | 1598.4
80 | 1624.1 | 1644.8
81 | 1637.2 | 1668.3
82 | 1643.6 | 1680.4
83 | 1638.6 | 1674.5
84 | 1621.5 | 1682.7
85 | 1635.9 | 1669.1
86 | 1657.6 | 1668
87 | 1650.2 | 1673.2
88 | 1648.2 | 1662.2
89 | 1640.4 | 1697.6
90 | 1671.5 | 1699.2
91 | 1661.8 | 1677.7
92 | 1638.6 | 1690.1
93 | 1655 | 1682.4
94 | 1646.9 | 1697.5
95 | 1636.2 | 1697.1
96 | 1645.4 | 1682.2
97 | 1657 | 1665.8
98 | 1661.1 | 1676
99 | 1647.2 | 1703.5
100 | 1649.9 | 1696.7
101 | 1647 | 1681.3
102 | 1653.5 | 1697.7
103 | 1622.1 | 1678.6
104 | 1647.5 | 1684.4
105 | 1659 | 1680.7
106 | 1620.3 | 1678.3
107 | 1635.6 | 1690.1
108 | 1631.1 | 1691.9
109 | 1641.8 | 1677.3
110 | 1642 | 1672.1
111 | 1642.3 | 1672.1
112 | 1645.9 | 1659.2
113 | 1641.5 | 1666.9
114 | 1651.8 | 1671
115 | 1635.4 | 1689.5
116 | 1627.3 | 1700.3
117 | 1626.5 | 1694.3
118 | 1625.2 | 1683.7
119 | 1626.4 | 1668.9
Baseline | 1632.3323 | 1671.1823
Minimum | 1259.9 | 1271.8
Half Minimum | 1446.1161 | 1471.4911
Instructions:
1. Copy the values from BC_SAG_SLICE_WIDTH_1.csv into cells B12:B131
2. Copy the values from BC_SAG_SLICE_WIDTH_2.csv into cells C12:C131
3. Check the Baseline shown in the Profile 1 chart. If necessary adjust the position, using the offset in cell B2
4. Check the Minimum shown in the Profile 1 chart . If necessary adjust the position, using the offset in cell B3
5. Check the Baseline shown in the Profile 2 chart. If necessary adjust the position, using the offset in cell B4
6. Check the Minimum shown in the Profile 2 chart . If necessary adjust the position, using the offset in cell B5
7. Copy BC_SAG_SLICE_WIDTH_ROIs.png over cell N11

## Cor Slice Width
Fitting Parameters | Value
Manual offset for baseline of profile 1 | 10
Manual offset for minimum of profile 1 | 30
Manual offset for baseline of profile 2 | 30
Manual offset for minimum of profile 2 | 10
Calculated Parameters | Value
FWHM profile 1 (pixels) | 26
FWHM profile 2 (pixels) | 26 |  | BC_COR_SLICE_WIDTH_ROIs
Pixel | Profile 1 | Profile 2
0 | 1701.6 | 1692.5
1 | 1710.5 | 1695.2
2 | 1727.4 | 1719.1
3 | 1742.2 | 1705.9
4 | 1716.5 | 1725
5 | 1702.9 | 1685.7
6 | 1746.3 | 1724
7 | 1733.4 | 1714.7
8 | 1736.6 | 1704.1
9 | 1713.7 | 1718.4
10 | 1733.2 | 1713.8
11 | 1757.2 | 1713
12 | 1756.5 | 1741.7
13 | 1732 | 1720
14 | 1727.1 | 1741.2
15 | 1744.6 | 1730.8
16 | 1747.8 | 1713.1
17 | 1744.6 | 1728.8
18 | 1761.2 | 1755
19 | 1768.9 | 1768.2
20 | 1752.6 | 1767.1
21 | 1758.4 | 1767.7
22 | 1787 | 1773.9
23 | 1776.1 | 1752.4
24 | 1782.9 | 1745.6
25 | 1767.6 | 1765.7
26 | 1760.7 | 1770.4
27 | 1775.7 | 1768.7
28 | 1775.3 | 1778.4
29 | 1760.8 | 1772.9
30 | 1791.7 | 1764.8
31 | 1772.2 | 1788.5
32 | 1752.5 | 1788.6
33 | 1751.2 | 1774.2
34 | 1726.2 | 1791.9
35 | 1704.6 | 1778.8
36 | 1677.4 | 1767.9
37 | 1668.9 | 1801.9
38 | 1624.9 | 1768.1
39 | 1586.7 | 1782.7
40 | 1527.9 | 1781.2
41 | 1508.5 | 1792.5
42 | 1461.4 | 1798.8
43 | 1428.4 | 1781
44 | 1426.2 | 1792.3
45 | 1392.1 | 1780.2
46 | 1382.4 | 1776.4
47 | 1360.4 | 1778.4
48 | 1391.4 | 1756.8
49 | 1366.7 | 1751.9
50 | 1363.7 | 1745
51 | 1367 | 1700.9
52 | 1374.1 | 1679.9
53 | 1368.2 | 1636.3
54 | 1388 | 1594.7
55 | 1385.8 | 1556.2
56 | 1353 | 1484.2
57 | 1343.3 | 1491.7
58 | 1377.9 | 1435.9
59 | 1383.3 | 1415.7
60 | 1393.9 | 1429.5
61 | 1415.9 | 1387
62 | 1430.9 | 1360.9
63 | 1459.2 | 1365.5
64 | 1534.5 | 1355.1
65 | 1566.9 | 1359.3
66 | 1603.1 | 1358
67 | 1646.5 | 1366.4
68 | 1664.4 | 1349.5
69 | 1707 | 1381.6
70 | 1733 | 1375
71 | 1742.8 | 1363.8
72 | 1797 | 1370.7
73 | 1793.1 | 1379.4
74 | 1781.6 | 1396.3
75 | 1810.1 | 1428.1
76 | 1776.9 | 1440.1
77 | 1791.3 | 1464.3
78 | 1791.1 | 1471.2
79 | 1802.8 | 1525.4
80 | 1787.4 | 1567.5
81 | 1804.3 | 1624.8
82 | 1801.5 | 1654.2
83 | 1789.3 | 1698.1
84 | 1805.5 | 1685
85 | 1806.9 | 1716.5
86 | 1797.3 | 1749.6
87 | 1805.1 | 1754.2
88 | 1777.2 | 1757
89 | 1770.5 | 1799.3
90 | 1794.9 | 1806.2
91 | 1802.6 | 1779.1
92 | 1802.8 | 1777
93 | 1782 | 1792.6
94 | 1803.5 | 1785.1
95 | 1808.4 | 1787.9
96 | 1794.2 | 1785.9
97 | 1797.4 | 1771.2
98 | 1782.9 | 1777.1
99 | 1780.4 | 1783.7
100 | 1781.6 | 1779.8
101 | 1797.9 | 1768.6
102 | 1783.9 | 1787.5
103 | 1766.9 | 1776.4
104 | 1784.9 | 1776.8
105 | 1778.6 | 1779.9
106 | 1771.6 | 1755.8
107 | 1788.9 | 1746.6
108 | 1779.4 | 1753.7
109 | 1777.8 | 1768.5
110 | 1768.2 | 1788
111 | 1768.6 | 1784.5
112 | 1778 | 1769.2
113 | 1756.9 | 1760.8
114 | 1776.6 | 1755.7
115 | 1753.3 | 1757.1
116 | 1761.7 | 1764.1
117 | 1774.2 | 1750.5
118 | 1755 | 1766.1
119 | 1751.2 | 1777.3
Baseline | 1773.9968 | 1785.6419
Minimum | 1373.3 | 1359.5
Half Minimum | 1573.6484 | 1572.571
Instructions:
1. Copy the values from BC_COR_SLICE_WIDTH_1.csv into cells B12:B131
2. Copy the values from BC_COR_SLICE_WIDTH_2.csv into cells C12:C131
3. Check the Baseline shown in the Profile 1 chart. If necessary adjust the position, using the offset in cell B2
4. Check the Minimum shown in the Profile 1 chart . If necessary adjust the position, using the offset in cell B3
5. Check the Baseline shown in the Profile 2 chart. If necessary adjust the position, using the offset in cell B4
6. Check the Minimum shown in the Profile 2 chart . If necessary adjust the position, using the offset in cell B5
7. Copy BC_COR_SLICE_WIDTH_ROIs.png over cell N11

## Slice Position
Slice Position
BC_TRA_SLICE_POSITION_LINES
Mean error + Std Dev Error (mm) | Mean Error (mm) | Mean error - Std Dev Error (mm) | Maximum (absolute) Error (mm)
0.33054 | -0.0019678 | -0.33447 | 0.63113
Input Parameters | Values
Slice thickness (mm) | 5
Slice gap (mm) | 2.5
Angled rod separation (mm) | 8
Distance between rods (mm) | 120
Calculated Parameters | Values
Slice offset (mm) | -104.9646
Slice | Distance (mm) | Mean Test Object Distance (mm) | Correction Factor | Sign | Slice Position (measured) | Slice Position (calculated) | Slice Position Error (mm) | Absolute Error
1 |  |  |  | -1
2 |  |  |  | -1
3 |  |  |  | -1
4 |  |  |  | -1
5 | 134.612 | 120.267 | 0.99778 | -1 | -67.0373 | -67.4646 | -0.42723 | 0.42723
6 | 119.627 | 119.768 | 1.0019 | -1 | -59.7957 | -59.9646 | -0.16885 | 0.16885
7 | 104.203 | 120.263 | 0.99781 | -1 | -51.8334 | -52.4646 | -0.63113 | 0.63113
8 | 90.209 | 119.812 | 1.0016 | -1 | -44.9978 | -44.9646 | 0.033259 | 0.033259
9 | 75.091 | 119.935 | 1.0005 | -1 | -37.3523 | -37.4646 | -0.1123 | 0.1123
10 | 60.553 | 120.021 | 0.99983 | -1 | -30.0058 | -29.9646 | 0.041182 | 0.041182
11 | 45.567 | 119.873 | 1.0011 | -1 | -22.4541 | -22.4646 | -0.01044 | 0.01044
12 | 30.653 | 119.874 | 1.0011 | -1 | -14.812 | -14.9646 | -0.15257 | 0.15257
13 | 16.776 | 119.827 | 1.0014 | -1 | -7.3866 | -7.4646 | -0.077979 | 0.077979
14 | 7.931 | 119.658 | 1.0029 | 1
15 | 17 | 120.017 | 0.99986 | 1 | 7.4986 | 7.5354 | 0.036786 | 0.036786
16 | 30.083 | 119.969 | 1.0003 | 1 | 14.5039 | 15.0354 | 0.5315 | 0.5315
17 | 45.929 | 120.017 | 0.99986 | 1 | 22.6101 | 22.5354 | -0.074729 | 0.074729
18 | 59.971 | 120.76 | 0.99371 | 1 | 29.5271 | 30.0354 | 0.50834 | 0.50834
19 | 75.538 | 119.917 | 1.0007 | 1 | 37.5829 | 37.5354 | -0.047457 | 0.047457
20 | 90.586 | 120.019 | 0.99984 | 1 | 45.1088 | 45.0354 | -0.073407 | 0.073407
21 | 105.973 | 119.769 | 1.0019 | 1 | 52.9378 | 52.5354 | -0.40237 | 0.40237
22 | 119.419 | 119.893 | 1.0009 | 1 | 59.6288 | 60.0354 | 0.40665 | 0.40665
23 | 134.016 | 119.89 | 1.0009 | 1 | 66.9501 | 67.5354 | 0.58533 | 0.58533
24 |  |  |  | 1
25 |  |  |  | 1
26 |  |  |  | 1
27 |  |  |  | 1
[6 empty] | Mean | -0.0019678
[6 empty] | Standard Deviation | 0.33251
[6 empty] | Maximum |  | 0.63113
Note on B13: Wastling,Stephen: This is the formula for calculating the intercept of a straight-line fit via linear least squares
Instructions:
1. Enter the slice thickness into cell B7
2. Enter the slice gap in cell B8
3. Copy the values from BC_TRA_SLICE_POSITION.csv into cells B17:C43
4. Copy BC_TRA_SLICE_POSITION_LINES.png over cell K3
5. Copy the Slice Position Error chart into the report

## EPI Stability
EPI Stability
Max signal | Min signal | Delta | Mean signal | Delta/Mean
772.56 | 771.653 | 0.907 | 772.0878 | 0.1175%
Per Set of 25 images
| 1 to 25 | 26 to 50 | 51 to 75 | 76 to 100
Mean Signal | 772.0676 | 772.0906 | 772.1318 | 772.0612
Signal Standard Deviation | 0.21469 | 0.13952 | 0.18902 | 0.14989
HC_TRA_EPI_STABILITY_ROI
Image Set | Signal
Images 1-25 | 772.222
Images 1-25 | 772.031
Images 1-25 | 772.284
Images 1-25 | 771.689
Images 1-25 | 772.276
Images 1-25 | 772.227
Images 1-25 | 771.902
Images 1-25 | 771.76
Images 1-25 | 772.364
Images 1-25 | 772.076
Images 1-25 | 771.933
Images 1-25 | 771.818
Images 1-25 | 772.062
Images 1-25 | 771.907
Images 1-25 | 771.818
Images 1-25 | 772.044
Images 1-25 | 772.351
Images 1-25 | 771.973
Images 1-25 | 771.689
Images 1-25 | 772.324
Images 1-25 | 772.316
Images 1-25 | 772.258
Images 1-25 | 771.987
Images 1-25 | 772.262
Images 1-25 | 772.116
Images 26-50 | 772.067
Images 26-50 | 772.009
Images 26-50 | 772.08
Images 26-50 | 771.893
Images 26-50 | 772.053
Images 26-50 | 772.129
Images 26-50 | 772.067
Images 26-50 | 772.476
Images 26-50 | 771.947
Images 26-50 | 772.018
Images 26-50 | 772.2
Images 26-50 | 771.938
Images 26-50 | 772.173
Images 26-50 | 772.262
Images 26-50 | 772.236
Images 26-50 | 772.16
Images 26-50 | 772.076
Images 26-50 | 771.978
Images 26-50 | 771.907
Images 26-50 | 772.084
Images 26-50 | 771.933
Images 26-50 | 772.276
Images 26-50 | 772.111
Images 26-50 | 771.956
Images 26-50 | 772.236
Images 51-75 | 772.164
Images 51-75 | 772.209
Images 51-75 | 772.196
Images 51-75 | 771.804
Images 51-75 | 771.96
Images 51-75 | 772.084
Images 51-75 | 772.089
Images 51-75 | 772.142
Images 51-75 | 771.907
Images 51-75 | 771.831
Images 51-75 | 772.053
Images 51-75 | 772.173
Images 51-75 | 772.018
Images 51-75 | 772.289
Images 51-75 | 772.453
Images 51-75 | 772.44
Images 51-75 | 772.067
Images 51-75 | 771.818
Images 51-75 | 772.182
Images 51-75 | 772.191
Images 51-75 | 772.178
Images 51-75 | 772.04
Images 51-75 | 772.56
Images 51-75 | 772.244
Images 51-75 | 772.204
Images 76-100 | 772.093
Images 76-100 | 772.173
Images 76-100 | 772.209
Images 76-100 | 772.262
Images 76-100 | 771.907
Images 76-100 | 772.116
Images 76-100 | 771.933
Images 76-100 | 772.178
Images 76-100 | 771.991
Images 76-100 | 772.098
Images 76-100 | 772.138
Images 76-100 | 771.653
Images 76-100 | 772.169
Images 76-100 | 772.333
Images 76-100 | 771.907
Images 76-100 | 772.164
Images 76-100 | 771.942
Images 76-100 | 772.071
Images 76-100 | 771.889
Images 76-100 | 772.147
Images 76-100 | 772.018
Images 76-100 | 771.978
Images 76-100 | 771.951
Images 76-100 | 772.231
Images 76-100 | 771.978
This chart isn’t available in your version of Excel. Editing this shape or saving this workbook into a different file format will permanently break the chart.
Instructions:
1. Copy the values from HC_TRA_EPI_STABILITY.csv into cells A12:A111
2. Copy HC_TRA_EPI_STABILITY_ROI.png over cell P11
3. Copy the Slice Variation Over Time into the report

## Tables for Report
Input Parameters | Value
System Name | Wellington South Aera 2025
Table 1: Body Coil NSNR | [5 empty] | Table 3: Body Coil Fractional Uniformity | [5 empty] | Table 5: Geometric Linearity for 3 Gradient Directions | [5 empty] | Table 7: Measured Slice Widths for a Nominal 5 mm Slice | [5 empty] | Table 8: Slice Position Error Measurements |  |  |  |  | Table 9: EPI Stability
Plane | Axial | Sagittal | Coronal | Mean |  | Gradient Direction | X | Y | Z | Mean | Standard Deviation | Gradient Direction | X | Y | Z | Mean | Standard Deviation | Plane | Axial | Sagittal | Coronal | Mean | Standard Deviation |  | Mean Error + Std Dev Error (mm) | Mean Error (mm) | Mean Error - Std Dev Error (mm) | Maximum (absolute) Error (mm) |  | Max signal | Min signal | Delta | Mean signal | Delta/Mean
Wellington South Aera 2025 | 448.8578 | 456.4287 | 471.7549 | 459.0138 |  | Wellington South Aera 2025 | 1 | 0.97813 | 0.90938 | 0.9625 | 0.043839 | Wellington South Aera 2025 | 120.137 | 121.0235 | 120.5135 | 120.558 | 0.39298 | Wellington South Aera 2025 | 5.1953 | 5.1953 | 5.1953 | 5.1953 | 0 | Wellington South Aera 2025 | 0.33054 | -0.0019678 | -0.33447 | 0.63113 | Wellington South Aera 2025 | 772.56 | 771.653 | 0.907 | 772.0878 | 0.1175%
Table 2: Head Coil NSNR | [5 empty] | Table 4: Head Coil Fractional Uniformity | [5 empty] | Table 6: Geometric Distortion (%) for 3 Gradient Directions
Plane | Axial | Sagittal | Coronal | Mean |  | Gradient Direction | X | Y | Z | Mean | Standard Deviation | Gradient Direction | X | Y | Z | Mean
Wellington South Aera 2025 | 3548.2066 | 2617.9988 | 2914.9089 | 3027.0381 |  | Wellington South Aera 2025 | 0.49062 | 0.38438 | 0.49375 | 0.45625 | 0.16752 | Wellington South Aera 2025 | 0.14356 | 0.058183 | 0.20764 | 0.13646
Delta/Mean
Instructions:
1. Fill in the name of the system being tested and the year of test in cell B2
2. Copy the completed tables into the report.

## Comparison Charts for Report
Input Parameters | Value
Comparison System Name | Wellington South Aera 2023
Body Coil NSNR
Wellington South Aera 2025 | 459.0138
Wellington South Aera 2023 | 430
Head Coil NSNR
Wellington South Aera 2025 | 3027.0381
Wellington South Aera 2023 | 3091
Body Coil Mean Uniformity
Wellington South Aera 2025 | 0.9625
Wellington South Aera 2023 | 0.98
Head Coil Mean Uniformity
Wellington South Aera 2025 | 0.45625
Wellington South Aera 2023 | 0.44
Mean Geometric Linearity Error (mm)
Wellington South Aera 2025 | 0.558
Wellington South Aera 2023 | 0.4
Mean Geometric Distortion (%)
Wellington South Aera 2025 | 0.13646
Wellington South Aera 2023 | 0.4
Slice Width Error (mm)
Wellington South Aera 2025 | 0.19531
Wellington South Aera 2023 | 0.21
Instructions:
1. Fill in the name of the system being compared with and the year of the comparison test in cell B2
2. Fill the comparison values into cells B8, B12, B16, B20, B24, B28 and B32
3. Copy the completed Charts into the report.